Spectrum Analyzer - Firmware version 2.78

1. Preliminary remarks

This firmware update (version 2.78) can be retrofitted in all spectrum analyzers. The firmware comprises all functions necessary for operation of the models FSA, FSAS, FSAC, FSAD, FSB, FSBS, FSBC, FSM, FSMS, FSMS26, ESAI, ESBI and ESMI. The configuration for the respective model is automatically displayed when the instrument is switched on.

After switching on, the instrument logs on with the following version numbers:

MP	V2.78
	GP	V2.70
	FP	V2.75

2. Hardware modifications

The Mainprocessor board 805.0012 in the Display Unit must be changed from version 02 to version 03 in the case of models FSA, FSAS, FSAC, FSAD, FSB, FSBS and FSBC.

In the case of models ESAI, ESBI, ESMI and FSM, version 03 is factory-fitted. Instruments with a firmware version of 2.01 and above are also updated.

For changing the EMI Test Receiver ESAI from version 2.01 to version 2.78 the EMI Board (1004.7006) must be changed (if a version > 2.01 is fitted, this has already been done).

Board	Modification	Reason


805.0012	R37 not fitted	Var02 -> Var03
Mainprocessor	R118 not fitted	for firmware update
	R119 not fitted	V1.XX -> V2.XX
	R121 fitted	FSA, FSAS, FSAC, FSAD
	Rewiring	FSB, FSBS, FSBC
	D10/4 - D1/27	(memory extension
	Rewiring	Paged Eproms)
	D84/27 - D58/27


1004.7006	R141 becomes 825 W	for firmware update
EMI detector		V2.01 -> V2.78

3.   Modifications and extensions of version V2.33

3.1 Modifications for all models

-	Normally, the printer is no longer reset after printing a hardcopy (due to respective control characters). A few printers (e.g. HP Deskjet 500) used to perform a formfeed because of this reset command. As a result of this, the ”Formfeed Off" function nevertheless caused a formfeed.

-	Support of the Option FSA-B7 (fast A/D converter) with AF Unit (500 kHz spurious FM DECT)   (V2.67 and above) with gated sweep (frequency domain), gapped sweep (time domain) with pretrigger and posttrigger

IEC-bus command:

Command              Data        Units            Meaning                                                                                        ERRORS?                                               Output error number                                                                                                                                         52 numeric data out                                                              of range (GAP)                                                                                                                                          Gap (only span = 0)       TRIGGER:GAP               ON                             switch on                                      OFF                            switch off                                                                                               :GAP:PRE TRIGGER    0s...100s       S          Change pretrig. time              :PRE TRIGGER?                   US                                                            INCREMENT       MS            increment                                       DECREMENT       NS            decrement                                                                                                 :GAP:POST TRIGGER   0s...100s       S          Change posttrg. time             :POST TRIGGER?                    US                                                            INCREMENT       MS            increment                                       DECREMENT       NS            decrement                                                                                                  :GAP:TIME           0s...100s      S          Change gap time                    :TIME?                          US                                                            INCREMENT       MS            increment                                       DECREMENT       NS            decrement                                                                                                                                                                                                                              Gate (only span <> 0)     TRIGGER:GATE              ON                             switch on                                      OFF                            switch off                                                                                                :GATE:DELAY        0s...100s       S          Change gate                           :DELAY?                       US         delay                                              INCREMENT       MS            increment                                       DECREMENT       NS            decrement                                                                                                  :GATE:LENGTH       0s...100s       S          Change gate length                    :LENGTH?                      US                                                            INCREMENT       MS            incmement                                       DECREMENT       NS            decrement                                                                                                  :GATE:LEVEL        0...63                     Gate threshold                        :LEVEL?                                                                                     INCREMENT                     increment                                       DECREMENT                     decrement                                                                                                                                                Gate mode:                      :GATE:MODE         EDGE                          edge triggering                                 LEVEL                         level triggering                                POSITIVE                      pos. edge/level                                 NEGATIVE                      neg. edge/level

DEMODULATION:FM             500kHz             HZ       (only Option FSA-B7)                                  50kHz             KHZ      (only Option FSA-B7)                                   5kHz             MHZ      (only Option FSA-B7)                                 500 Hz             GHZ      (only Option FSA-B7)

-	Quasi-analog display of sweep for trace 1 and trace 2

IEC-bus commands:

          Command              Data        Units            Meaning

Analog trace display                                                          in trace 1    DISPLAY:ANALOG:T_1        ON                           - on                                          OFF                          - off                                                                    Analog trace display                                                         in trace 2                          :T_2        ON                           - on                                           OFF                         - off

-	Special function YIGSYN NORMAL/NARROW for optimizing the phase noise in the zero span

IEC-bus commands:

Command              Data        Units            Meaning

Setting YIG-SYN:     YIG SYN                   NORMAL                     - normal                                           NARROW                     - always narrowband

-	In coupled mode, the sweep time can be limited to a minimum sweep time (matching the sweep time to the repetition rate of pulse-shaped input signals).

IEC-bus commands:

          Command               Data         Units            Meaning

SWEEP:MIN_TIME       20ms-1980s       S          Sweep time (coup.Mode)                                                                                 :MIN_TIME?                      US                                                              INCREMENT       MS            incmement                                         DECREMENT       NS            decrement

ON                          - on                                                                       OFF                         - off

-	List Mode. The values of the display and frequency lines as well as the difference between them are additionally indicated on the screen.

IEC-bus commands:

          Command               Data         Units            Meaning

List Screen    DISPLAY:LIST                ON                           - on                                            OFF                         - off

-    Measurement of occupied bandwidth and adjacent channel power'

IEC-bus commands:

CHANNEL:ACP?                   --            --      Adjacent Channel Power          :OBW?                   --            --      Occupied Bandwidth              :BANDWIDTH           0...max          HZ      Channel Bandwidth               :BANDWIDTH?             --           KHZ                                                                           MHZ                                                                           GHZ                                      :SPACE               0...max          HZ      Channel Space                   :SPACE?                 --           KHZ                                                                           MHZ                                                                           GHZ

-    The message UNCAL can be cut off via a service function.

Service function  0.99.81

-    New service function Yig-filter calibration

Service function  0.05.01

3.2 Modifications for ESAI, ESBI:

-	In the case of V2.33 and below, a digital Max.Hold was performed in the measurements during the set measuring time in scan mode with AVG and RMS detector. In the case of versions > V2.33 the average of all measurements is taken. This will result in a lower displayed value if the pulse repetition rate of the input signal is smaller than the averaging time constant of the respective detector.

-	In Scan Mode a log step size is permissible even without Config Table (operation in the sweep menu, the Partial Scan function has been shifted to the right-hand supplementary menu).

IEC-bus commands:

SCAN:STEPSIZE	value	HZ	(lin)	Step size in Hz
SCAN:STEPSIZE?		KHZ	(lin)
		MHZ	(lin)
		PCT	(log.)	Step size in %
	INCREMENT
	DECREMENT
	LINEAR			Lin.  step size
	LOGARITHMIC			Log. step size

-	When changing to the EMI Receiver Mode and with the Probe Code connected, the unit is switched to dBuV unless the Probe Code specifies a unit. In earlier versions, the unit remained dBm in this case.

3.3 Modifications for FSM (see Supplements to Manual):

-	new function External Mixer (see Supplements to Manual) with mixer bias setting, selection of harmonic number, signal ID functions, input 1st LO range.

IEC-bus commands:

Command              Data           Units        Meaning                                                                                          AVG CONV LOSS            -116...+120dB      DB        Offset reference        AVG CONV LOSS?                                         level                                              INCREMENT        --                                                        DECREMENT        --                                                                                                             HARMONIC:NUMBER           3...200          --        Selection of                                                            desired                                                                                         harmonic                                      INCREMENT        --                                                        DECREMENT        --                                                                                                                                                                 Frequ. range of                                                           1st Local Osc.     LO RANGE:DEFAULT           --               --         Default                                                                                                       :MIN FREQ        5.27...12.14GHz     HZ        Lower frequency                                                              limit                 :MIN FREQ?                          KHZ                                                          INCREMENT         MHZ                                                          DECREMENT         GHZ                                                                                                                      :MAX FREQ        6.27...13.14GHz     HZ        Upper frequency                                                                             limit                       :MAX FREQ?                          KHZ                                                          INCREMENT         MHZ                                                          DECREMENT         GHZ                                                                                                                     MIXER            INTERNAL           --         Ext. mixer on                                  EXTERNAL          --         Ext. Mixer off             :BIAS              -2.047V...         V         Settg. mixer bias             :BIAS?            +2.047V           MV                                                                         UV                                                                           NV                                                                                                                                                                      Setting the bias                                            ON              --        active                                                   OFF             --        inactive                                                                                  SIGNAL ID                 ON             --       Signal ident. on                                     OFF            --       Signal ident off                                                                                          :CALCULATE                                     Filter mixer                                                                   products                     :CONTINUE                                     Continue signal                                                            identification

3.4 New Modell ESMI (V2.60 and above)

3.5 Normalization with tracking generator:

In the case of version 2.75 and above, the normalization is no longer switched off automatically when the LEVEL RANGE is changed. Changing the range with normalization activated causes a higher graphical resolution with the same dynamic range.

Note: The display range (LEVEL RANGE) used for calibration of the transmission or reflection measurement determines the useful dynamic range.

If, with normalization activated, the display range is increased compared to the calibration setting, it is advisable to perform another calibration.

3.6 Support of Models FSMS26 and ESMI with Option TG up to 26.5 GHz (V2.70 and above)

4.   Error corrections

4.1 Errors eliminated for V2.50 and above

-	ESAI, ESBI: In Scan Mode with LOG step size and stop frequencies > 1 GHz, a numerical overflow occurs in the frequency calculation. As a result of this, the frequency stepping to the next test point fails, and the final frequency of the scan is not reached (error for versions below V2.33).

-	ESAI, ESBI: When using transducer tables by means of special function (without Range Table), an error occurs in the calculation of the correction value. In Analyzer and Overview mode, a correction value of 0 dB is assumed for frequencies lying outside the transducer table. In Scan mode, the lower or higher value of the table is used as correction value in this case (error in V2.33).

- 	ESAI, ESBI: In Scan mode with Autorange On and Attenuation Manual 0 dB, the autorange function does not work properly (takes 3 times longer than with Att. 10 dB (error for versions below V2.33). Remedy: Set Autorange Min Att. to 0 dB.

-	FSM/FSMS: occasional error message YIG Fine Tuning after control via IEC bus with
"*RST; F:C 100MHZ;SPAN 0" (error  < V2.50)

-	Generally: After writing to the Plot text field via IEC bus it may happen that the text is not output on the printer (error for versions < V2.50).

-	ESAI, ESBI: In Scan mode with the Config Table activated (with bandwidth switching), the measurement is made too early with the RMS detector (without input signal, only inherent noise) after switching over the bandwidth. After switching over the bandwidth, the measured data feature a transient response decreasing only gradually (error for versions < V2.50).

-	ESAI, ESBI: The settling time for the EMI bandwidths 100 Hz and 200 Hz has been prolonged because the measurement was sometimes made too early (spike occurs at the left edge of the screen due to return from the stop to the start frequency)  (error for versions  < V2.50).

-	Generally: In log. Span a sinewave signal lying directly at the turning points of the linear partial sweeps may show a wrong level or be even totally suppressed (error for versions  < V2.50).

-	ESAI, ESBI: Scan Mode with activated Config Table and step size LOG 100 %. The actual step size is considerably smaller than the step size entered (eliminated for MP V2.50 and above). Remedy: Enter step size 99.99%.

4.2 Errors eliminated for V2.51 and above

-	Correction of the interpolation of the transducer tables at the right edge of the table. Overshoots may occur if there are great changes between the last two sampling points.

-	In the case of version FP V2.50 the frequency response correction does not work properly with external mixing, which may result in level errors of a few dB (eliminated for FP V2.51 and above).

4.3 Errors eliminated for V2.64 and above

-	ESBI: In the case of versions FP V2.50 to FP V2.61 the control of the preselection filters does not work properly with RF preselection activated. As a result of this, wrong filter ranges are selected or no preselection filter is cut in at all. This results in sudden level changes by several dB at the edges where the filters switch, because the frequency response correction uses the correct values (eliminated for FP V2.62 and above). This error does not occur in version V2.40.

-	FSM, ESMI: In version FP V2.62 level errors occur with small sweep times at the beginning of the sweep in the case of spans from 5 MHz to 50 MHz (for center frequency < 12.9 GHz) or spans from 10 MHz to 50 GHz (for center frequency >12.9 (eliminated for version 2.64 and above).

-	In the case of versions GPF V2.50 to GPF V2.60 a level difference of 0.3 to 0.5 dB occurs in Analyzer Mode between the detectors AUTOPEAK and POS. This error is caused by a faulty control of the detector selection (eliminated for GPF V2.63 and above). This error does not occur in conjunction with the Option FSA-B7 (Fast A/D converter)

-	All models: Log Span in Analyzer or Overview Mode (for ESAI, ESBI, ESMI). When combining measured values of the linear partial sweeps, signals lying exactly at the frequency of the thresholds between the partial intervals can be suppressed or attenuated (eliminated for MP V2.60 and above).

-	All models: Log Span in Analyzer or Overview Mode (for ESAI, ESBI, ESMI). Due to an error in the combination of measured values, phantom signals may be indicated. They appear as 1-pixel wide peaks irrespective of the resolution bandwidth set (eliminated for MP V2.64 and above).

-	All models with Option FSA-B7 (fast A/D converter): The short calibration CAL SHORT does not work properly if the span before calling CAL SHORT > 0 (eliminated for MP V2.64 and above).

-	ESBI: Control of RF preselection does not work properly in the case of versions V2.45 to V2.60, however switchover of the frequency response correction works properly so that level jumps of several dB may occur in the individual filter ranges (eliminated for V2.64 and above)

-	FSx, ESxI with Option FSA-B7 (fast A/D converter): CAL SHORT does not work properly (eliminated for V2.63 and above)

-	FSx, ESxI: In the case of logarithmic span, a signal may additionally occur as 1-pixel wide line (irrespective of the set bandwidth) due to an error in the combination of measured values of the linear partial span. This phantom signal always features a lower level than the real signal (error in versions V2.50, V2.60, eliminated for V2.64 and above)

4.4 Errors eliminated for V2.67 and above

-	FSx, ESxI: Transducer input 0.091 MHz results in the display 9.09999 MHz (error eliminated for V2.66 and above)

-	FSM, ESAI, ESBI, ESMI: Frequency response programming does not function properly in version V2.64, remedy is possible by matching the programming routine (error eliminated for V2.67 and above)

4.5 Errors eliminated for V2.78 and above

-	All models: If ACTIVE WINDOW 2 is selected in Split Screen, the display and labelling of the active grid is not correct after switching off/on. Remedy: Switch to ACTIVE WINDOW 1 before switching off. (error eliminated for 2.74 and above)

-	UNCAL is not indicated properly. Error for V2.70 and above  (error eliminated for 2.74 and above)

-	Error when forming the difference between Trace1-Trace2. Error for V2.75 (error eliminated for 2.76 and above)

-	Faulty values are used in the transducer input. Error in the last digit. (error eliminated for V2.75 and above)

-	ESMx: Quasi-peak indication fluctuates by 2 dB in the case of certain pulse sequences (error eliminated for V2.75 and above)

-	Yig-Sync error if marker count is activated in the multi-range sweep (error eliminated for V2.75 and above)

-	After restart, the sweep is shifted by 1 pixel compared with the subsequent sweeps (error eliminated for V2.70 and above)

-	Filter calibration for models FSM../ESM.. with TG up to 5 GHz not inhibited over 5 GHz. (error eliminated for V2.70 and above)

-	Instrument stops with LOG sweep and normalization. (error eliminated for V2.75 and above)

-	In the case of shape factor measurements on broadband filters, the output is wrong. (error eliminated for V2.75 and above)

-	IEC-bus  command TCT may cause problems in the handshake (error for 2.78 and above)

-	Stop frequency of 26.5 GHz cannot be set via the numerical keypad or the IEC bus. (error eliminated for 2.78 and above)

4.6 Error in the manual

-	ESxI  The following command is not correct in the IEC-bus command list:

DEMODULATION:AM PULSE

correct:

DEMODULATION:AM_PULSE

Models and Firmware versions

|  | Model | starting from version |
| --- | --- | --- |
|  | FSA | 1.30 |
|  | FSAS | 1.30 |
|  | FSAC | 1.30 |
|  | FSAD | 1.50 |
|  | FSB | 1.40 |
|  | FSBS | 1.40 |
|  | FSBC | 2.20 |
|  | FSM | 2.01 |
|  | FSM (33GHz) | 2.33 |
|  | FSMS (TG 5GHz) | 2.40 |
|  | FSMS (TG 26.5 GHz) | 2.70 better 2.74 |
|  | FSM External mixer | 2.50 |
|  | ESAI | 2.01 |
|  | ESBI | 2.10 |
|  | ESMI | 2.60 |
|  | ESMI with ESMI-B1 (TG up to 26.5 GHz) | 2.70 better 2.74 |
|  | FS-K1 MSDOS 3.1 | 1.30 |
|  | FS-K1 MSDOS 3.2 | 1.30 |
|  | FSA-B7 (Fast A/D converter) | 2.50 |
|  | Var 55 (DECT demodulator + FSA-B7) | 2.67 |

Caution: Not suitable for ESBI: > V2.40 incl. V2.60 because of error in preselection switching !!

5.   Calibrating the Level Accuracy of the RF Filters

µ §

If a tracking generator is not fitted when the CALIBRATE FILTER function is called, the calibration can't be performed. In this case, the CALIBRATE FILTER softkey has no function.