�                 �      �������͸                 �

�                 �          �                   �����

�       �����ͻ   �����ͻ    �    �����ͻ �����͸  �     ��Ŀ   ��Ŀ  Ŀ

�       �     �   �     �    �    �����ͼ �����ͻ  �     ����   ����   �

�����; �������; �����ͼ    �    �����; �����ͼ  ��;  ���� � ����  ���

LabTest 2.21 instrument file documentation for

Rohde & Schwarz Signal Generator SMH

July 1991

_______________________________________________________________________________

Introduction:

This instrument module provides GPIB programming support for the Rohde &

Schwarz Signal Generator SMH. It contains functions for opening,

configuring, and closing the instrument via the GPIB.

_______________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

- the SMH is connected to the GPIB.

- the  GPIB address supplied to the Initialize function must match

the GPIB address set on the SMH

- there is one unused device name in the handler.

_______________________________________________________________________________

LabWindows Libraries:

In  order to use this instrument module outside of the LabWindows envi-

ronment, it must be linked with the following LabWindows libraries:

- Formatting and I/O

- GPIB

_______________________________________________________________________________

Error Codes:

All error codes are returned in a global integer variable whose name

depends on the language being used:

For QuickBASIC : smh.err%

For C          : smh_err

A program should examine this variable after each function call is made

to determine if any errors occurred.

________________________________________________________________________________

Function Tree Layout:

Initialize

Freq./Level Setting

Modulation Setting

RF Sweep Setting

AF Sweep Setting

Basic Setting...

RF Setting

RF Level Setting

AF Setting

AF Level Setting

Freq./Level Offset

Special Setting

Variation Settings...

Freq./Level Steps

Modulation Steps

Special Functions...

Single Variation Step

Operation Complete

Diagnostics

Close

Driver Version

Get Handle

Set Operation Complete Query

________________________________________________________________________________

Function           : Initialize

QuickBASIC Syntax  : dy% = smh.init% (dev%, addr%)

C Syntax           : int smh_init (dev, addr);

int dev, addr;

Purpose            : Initializes the SMH on the GPIB and resets the

instrument to its power-on default state.

Parameter List     : dev  - SMH ID number (0 to 3)

addr - GPIB Address (0 to 30)

QuickBASIC Example : dy% = smh.init% (0, 16)

C Example          : int dy;

dy = smh_init (0, 16);

Error Codes        : The valid error codes for this function are:

0 : Success

-1 : GPIB Address out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

220 : Unable to open instrument

230 : GPIB write error

233 : Unable to configure GPIB address

236 : Unable to clear instrument

________________________________________________________________________________

Function           : Freq./Level Setting

QuickBASIC Syntax  : CALL smh.set.freq.level (dev%, rf#, rfun%, rflvl#,

rflun%, rfb%, af#, afun%,

aflvl%, afb%)

C Syntax           : void smh_set_freq_level (dev, rf, rfun, rflvl, rflun,

rfb, af, afun, aflvl, afb);

int dev;

double rf;

int rfun;

double rflvl;

int rflun, rfb;

double af;

int afun, aflvl, afb;

Purpose            : Configures the signal generators frequency and

level parameters for both RF and AF.

Parameter List     : dev  - SMH ID number (0..3)

rf - RF frequency

rfun  - RF frequency unit

0 : Hz

1 : kHz

2 : MHz

3 : GHz

rflvl - RF level

rflun - RF level unit

0 : dBm

1 : dB�V

2 : mV

3 : �V

rfb - RF level setting

0 : Off

1 : Normal

2 : EMF

af - AF frequency

afun - AF frequency unit

0 : Hz

1 : kHz

aflvl - AF level (mV)

afb - AF level setting

0 : Off

1 : On

QuickBASIC Example : CALL smh.set.freq.level (0,100.0,2,-30.0,0,1,1.0,1,1000,0)

C Example          : smh_set_freq_level (0,100.0,2,-30.0,0,1,1.0,1,1000,0);

Error Codes        : The valid error codes for this function are:

0 : Success

-2 : RF frequency unit out of range

-3 : RF level unit out of range

-4 : RF level setting out of range

-5 : AF frequency unit out of range

-6 : AF level setting out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

300 : One or more device dependent parameter(s)

were out of range within the previous

command(s).

________________________________________________________________________________

Function           : Modulation setting

QuickBASIC Syntax  : CALL smh.set.modulation (dev%, fp%, ap%, phm#, fm#, am#)

C Syntax           : void smh_set_modulation (dev, fp, ap, phm, fm, am);

int dev, fp, ap;

double phm, fm, am;

Purpose            : Configures the instrument's modulation parameters.

Parameter List     : dev  - SMH ID number (0..3)

fp - FM/PHM setting

0 : FM/PHM Off

1 : FM Int.

2 : FM Ext. AC

3 : FM Ext. DC

4 : FM Ext. AC + Int.

5 : FM Ext. DC + Int.

6 : FM FSK AC

7 : FM FSK DC

8 : PHM Int.

9 : PHM Ext. AC

10 : PHM Ext. AC + Int.

ap - AM/Pulse setting

0 : AM/Pulse Off

1 : AM Int.

2 : AM Ext. AC

3 : AM Ext. DC

4 : AM Ext. AC + Int.

5 : AM Ext. DC + Int.

6 : Pulse

phm - PHM (rad)

fm - FM (kHz)

am - AM (%)

QuickBASIC Example : CALL smh.set.modulation (0,1,1,1.0,10.0,30.0)

C Example          : smh_set_modulation (0,1,1,1.0,10.0,30.0);

Error Codes        : The valid error codes for this function are:

0 : Success

-8 : FM/PHM setting out of range

-9 : AM/Pulse setting out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

300 : One or more device dependent parameter(s)

were out of range within the previous

command(s).

________________________________________________________________________________

Function           : RF Sweep Setting

QuickBASIC Syntax  : CALL smh.set.rf.sweep (dev%, md%, tim#, endf#, ssiz#,

startf#, ssizun%, efun%, stfun%)

C Syntax           : void smh_set_rf_sweep (dev, md, tim, endf, ssiz,

startf, ssizun, efun, stfun);

int dev, md;

double tim, endf, ssiz, startf;

int ssizun, efun, stfun;

Purpose            : Configures the instrument for RF Sweep mode and sets

the corresponding parameters.

Parameter List     : dev  - SMH ID number (0..3)

md - Mode

0 : Auto

1 : Single

2 : Manual

3 : Reset

4 : Off

tim - Time per Step (ms)

endf - Stop frequency

ssiz - Step size

startf - Start frequency

ssizun - Step size unit

0 : Hz (linear steps)

1 : kHz (linear steps)

2 : MHz (linear steps)

3 : GHz (linear steps)

4 : % (log. steps)

efun - Stop frequency unit

0 : Hz

1 : kHz

2 : MHz

3 : GHz

stfun - Start frequency unit

0 : Hz

1 : kHz

2 : MHz

3 : GHz

QuickBASIC Example : CALL smh.set.rf.sweep (0,4,10.0,1000.0,1.0,1.0,2,2,2)

C Example          : smh_set_rf_sweep (0,4,10.0,1000.0,1.0,1.0,2,2,2);

Error Codes        : The valid error codes for this function are:

0 : Success

-13 : Step size unit out of range

-14 : Stop frequency unit out of range

-15 : Start frequency unit out of range

-16 : Mode out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

300 : One or more device dependent parameter(s)

were out of range within the previous

command(s).

________________________________________________________________________________

Function           : AF Sweep Setting

QuickBASIC Syntax  : CALL smh.set.af.sweep (dev%, md%, tim#, endf#, ssiz#,

startf#, ssizun%, efun%, stfun%)

C Syntax           : void smh_set_af_sweep (dev, md, tim, endf, ssiz,

startf, ssizun, efun, stfun);

int dev, md;

double tim, endf, ssiz, startf;

int ssizun, efun, stfun;

Purpose            : Configures the instrument for AF Sweep mode and sets

the corresponding parameters.

Parameter List     : dev  - SMH ID number (0..3)

md - Mode

0 : Auto

1 : Single

2 : Manual

3 : Reset

4 : Off

tim - Time per Step (ms)

endf - Stop frequency

ssiz - Step size

startf - Start frequency

ssizun - Step size unit

0 : Hz (linear steps)

1 : kHz (linear steps)

2 : % (log. steps)

efun - Stop frequency unit

0 : Hz

1 : kHz

stfun - Start frequency unit

0 : Hz

1 : kHz

QuickBASIC Example : CALL smh.set.af.sweep (0,4,10.0,100.0,1.0,1.0,1,1,1)

C Example          : smh_set_af_sweep (0,4,10.0,100.0,1.0,1.0,1,1,1);

Error Codes        : The valid error codes for this function are:

0 : Success

-17 : Step size unit out of range

-18 : Stop frequency unit out of range

-19 : Start frequency unit out of range

-20 : Mode out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

300 : One or more device dependent parameter(s)

were out of range within the previous

command(s).

________________________________________________________________________________

Function           : RF Setting

QuickBASIC Syntax  : CALL smh.set.rf (dev%, rf#, rfun%)

C Syntax           : void smh_set_rf (dev, rf, rfun);

int dev;

double rf;

int rfun;

Purpose            : Sets the signal generators RF frequency

Parameter List     : dev  - SMH ID number (0..3)

rf - RF frequency

rfun  - RF frequency unit

0 : Hz

1 : kHz

2 : MHz

3 : GHz

QuickBASIC Example : CALL smh.set.rf (0,100.0,2)

C Example          : smh_set_rf (0,100.0,2);

Error Codes        : The valid error codes for this function are:

0 : Success

-2 : RF frequency unit out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : RF Level Setting

QuickBASIC Syntax  : CALL smh.set.rf.level (dev%, rflvl#, rflun%, rfb)

C Syntax           : void smh_set_rf_level (dev, rflvl, rflun, rfb);

int dev;

double rflvl;

int rflun;

double rfb;

Purpose            : Set the signal generators RF level.

Parameter List     : dev  - SMH ID number (0..3)

rflvl - RF level

rflun - RF level unit

0 : dBm

1 : dB�V

2 : mV

3 : �V

rfb - RF level setting

0 : Off

1 : Normal

2 : EMF

QuickBASIC Example : CALL smh.set.rf.level (0,-30.0,0,1)

C Example          : smh_set_rf_level (0,-30.0,0,1);

Error Codes        : The valid error codes for this function are:

0 : Success

-3 : RF level unit out of range

-4 : RF level setting out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : AF Setting

QuickBASIC Syntax  : CALL smh.set.af (dev%, af#, afun%)

C Syntax           : void smh_set_af (dev, af, afun);

int dev;

double af;

int afun;

Purpose            : Sets the signal generators AF frequency.

Parameter List     : dev  - SMH ID number (0..3)

af - AF frequency

afun  - AF frequency unit

0 : Hz

1 : kHz

QuickBASIC Example : CALL smh.set.af (0,1.0,1)

C Example          : smh_set_af (0,1.0,1);

Error Codes        : The valid error codes for this function are:

0 : Success

-5 : AF frequency unit out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : AF Level Setting

QuickBASIC Syntax  : CALL smh.set.af.level (dev%, aflvl%, afb%)

C Syntax           : void smh_set_af_level (dev, aflvl, afb);

int dev, aflvl, afb;

Purpose            : Set the signal generators AF Level.

Parameter List     : dev  - SMH ID number (0..3)

aflvl - AF level (mV)

afb - AF level setting

0 : Off

1 : On

QuickBASIC Example : CALL smh.set.af.level (0,1000,0)

C Example          : smh_set_af_level (0,1000,0);

Error Codes        : The valid error codes for this function are:

0 : Success

-6 : AF level setting out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Freq./Level Offset

QuickBASIC Syntax  : CALL smh.set.freq.level.offset (dev%, lofs#, rofs#, rfun%)

C Syntax           : void smh_set_freq_level_offset (dev, lofs, rofs, rfun);

int dev;

double lofs, rofs;

int rfun;

Purpose            : Set the SMH's frequency and level offsets.

Parameter List     : dev  - SMH ID number (0..3)

lofs - Level offset (dB)

rofs - RF offset

rfun - RF offset unit

0 : Hz

1 : kHz

2 : MHz

QuickBASIC Example : CALL smh.set.freq.level.offset (0,0.0,0.0,2)

C Example          : smh_set_freq_level_offset (0,0.0,0.0,2);

Error Codes        : The valid error codes for this function are:

0 : Success

-7 : RF Offset unit setting out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Special settings

QuickBASIC Syntax  : CALL smh.special.setting (dev%, ref%, lcb%, lvls%)

C Syntax           : void smh_special_setting (dev, ref, lcb, lvls);

int dev, ref, lcb, lvls;

Purpose            : Configures special SMH functions.

Parameter List     : dev  - SMH ID number (0..3)

ref - Reference frequency

0 : External

1 : Internal

lcb - Level control

0 : Off

1 : On

lvls - Non-Interrupting level setting

0 : Off

1 : On

QuickBASIC Example : CALL smh.special.setting (0,1,1,0)

C Example          : smh_special_setting (0,1,1,0);

Error Codes        : The valid error codes for this function are:

0 : Success

-10 : Reference frequency setting out of range

-11 : Level control setting out of range

-12 : Non-Int. Level setting out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Freq./Level Steps

QuickBASIC Syntax  : CALL smh.set.freq.level.step (dev%, rfst#, rfun%,

rflvlst#, afst#, afun%,

aflvlst%)

C Syntax           : void smh_set_freq_level_step (dev, rfst, rfun, rflvlst,

afst, afun, aflvlst);

int dev;

double rfst;

int rfun;

double rflvlst, afst;

int afun, aflvlst;

Purpose            : Configures frequency and level steps for both

RF and AF.

Parameter List     : dev  - SMH ID number (0..3)

rfst - RF step

rfun - RF step unit

0 : Hz

1 : kHz

2 : MHz

3 : GHz

rflvlst - RF Level step (dB)

afst - AF step

afun - AF step unit

0 : Hz

1 : kHz

aflvlst - AF Level step (mV)

QuickBASIC Example : CALL smh.set.freq.level.step (0,1.0,2,1.0,1.0,1,1)

C Example          : smh_set_freq_level_step (0,1.0,2,1.0,1.0,1,1);

Error Codes        : The valid error codes for this function are:

0 : Success

-21 : AF step unit out of range

-22 : RF step unit out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Modulation Step

QuickBASIC Syntax  : CALL smh.set.modulation.step (dev%, am#, fm#, phm#)

C Syntax           : void smh_set_modulation_step (dev, am, fm, phm);

int dev;

double am, fm phm;

Purpose            : Configures SMH's modulation steps for AM, FM and PHM.

Parameter List     : dev  - SMH ID number (0..3)

am - AM step (%)

fm - FM step (kHz)

phm - PHM step (red)

QuickBASIC Example : CALL smh.set.modulation.step (0,1.0,1.0,1.0)

C Example          : smh_set_modulation_step (0,1.0,1.0,1.0);

Error Codes        : The valid error codes for this function are:

0 : Success

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Single Variation Step

QuickBASIC Syntax  : CALL smh.single.var.step (dev%, fun%, stp%)

C Syntax           : void smh_single_var_step (dev, fun, stp);

int dev, fun, stp;

Purpose            : Executes one single variation steps.

Parameter List     : dev  - SMH ID number (0..3)

fun - Function

0 : AF

1 : RF

2 : AF Level

3 : RF Level

4 : AM

5 : FM

6 : PHM

stp - Step

0 : Down

1 : Up

QuickBASIC Example : CALL smh.single.var.step (0,0,0)

C Example          : smh_single_var_step (0,0,0);

Error Codes        : The valid error codes for this function are:

0 : Success

-23 : Step out of range

-24 : Function out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Operation Complete

QuickBASIC Syntax  : CALL smh.op.complete (dev%, opc%)

C Syntax           : void smh_single_var_step (dev, opc);

int dev, *opc;

Purpose            : Waits for execution of all previous commands and

returns an error code if one or more device

dependent parameters were out of range.

Parameter List     : dev  - SMH ID number (0..3)

opc - Operation Complete

(Function returns when opc = 1)

QuickBASIC Example : CALL smh.op.complete (0,opc%)

C Example          : int opc;

smh_op_complete (0,&opc);

Error Codes        : The valid error codes for this function are:

0 : Success

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

300 : One or more device dependent parameter(s)

were out of range within the previous

command(s).

________________________________________________________________________________

Function           : Diagnostics

QuickBASIC Syntax  : CALL smh.diagnostics (dev%, testp%, test%, rdng#)

C Syntax           : void smh_single_var_step (dev, testp, test, rdng);

int dev, testp, test;

double *rdng;

Purpose            : Returns the voltage on one of the SMH's internal

testpoints.

Parameter List     : dev  - SMH ID number (0..3)

testp - Internal testpoint

test - Turns test on or off

0 : Off

1 : On

rdng - Voltage at the selected testpoint

QuickBASIC Example : CALL smh.diagnostics (0,0,1,rdng#)

C Example          : double rdng;

smh_diagnostics (0,0,1,&rdng);

Error Codes        : The valid error codes for this function are:

0 : Success

-25 : Test out of range

217 : Illegal SMH ID number

218 : Instrument not initialized

230 : GPIB write error

________________________________________________________________________________

Function           : Close

QuickBASIC Syntax  : CALL smh.close (dev%)

C Syntax           : void smh_close (dev);

int dev;

Purpose            : Effectively disconnects the SMH from the GPIB

and returns the instrument to local mode.

Parameter List     : dev  - SMH ID number (0..3)

QuickBASIC Example : CALL smh.close (0)

C Example          : smh_close (0);

Error Codes        : The valid error codes for this function are:

0 : Success

217 : Illegal SMH ID number

221 : Unable to close instrument

234 : Unable to place instrument in local mode

________________________________________________________________________________

Function           : Driver Version

QuickBASIC Syntax  : CALL smh.driver.version (v#)

C Syntax           : void smh_driver_version (v);

double *v;

Purpose            : Returns the actual version number of the SMH

instrument driver.

Parameter List     : v - Version number of SMH instrument driver

QuickBASIC Example : CALL smh.driver.version (version#)

C Example          : double version;

smh_driver_version (&version);

Error Codes        : Unused

________________________________________________________________________________

Function           : Get Handle

QuickBASIC Syntax  : hdl% = smh.get.handle% (dev%)

C Syntax           : int smh_get_handle (dev);

int dev;

Purpose            : Returns the GPIB Device handle of the specified

SMH device. This handle is needed by a programmer

to perform additional ib...() calls to the instrument

in order to implement functions which may be not

supported by the current version of the instrument

driver.

Parameter List     : dev - SMH ID number (0..3)

QuickBASIC Example : hdl% = smh.get.handle% (0)

CALL ibrsp (hdl%, spr%)

C Example          : int hdl, spr;

hdl = smh_get_handle (0);

ibrsp (hdl, &spr);

Error Codes        : The valid error codes for this function are:

0 : Success

217 : Illegal SMH ID number

218 : Instrument not initialized

________________________________________________________________________________

Function           : Set Operation Complete Query

QuickBASIC Syntax  : CALL smh.set.opc (onoff%)

C Syntax           : void smh_set_opc (onoff);

int onoff;

Purpose            : Enables Operation Complete Query in low-level

functions in order to test for execution errors

(i. e. parameter(s) out of range, etc...).

Disabling Operation Complete Query increases exe-

cution speed of low-level functions, but does not

change the operation of high-level functions, where

OPC Query is always enabled!

Parameter List     : onoff - OPC Query On or Off

<> 0: On

=  0: Off

QuickBASIC Example : CALL smh.set.opc (0)

C Example          : smh_set_opc (0);

Error Codes        : Unused

________________________________________________________________________________