INSTRUCTIONS ABOUT NEXUS SW UPDATE PROGRAM

==========================================

Note:

The attached Nexus application program is version 2.10.

When running the update program the version will appear as 02_10_00

(and the old version as 02_01_00).

In the Nexus the new version will appear as 2  1  0  (in selftest mode).

1) This is a DOS program.

To start the SW update type at the DOS prompt: a:update <Enter>.

When started follow the instructions. Note that Shorting Adaptor JP0918 (for shorting pins 2 and 3 of the RS232) is required.

(The program can also be started from Windows with proper settings.)

2) If an error message is displayed when the program try to read the actual

software version, you have probably chosen the wrong com port (even if

"communication ok" has been displayed). Then restart the program

and change the com port setting.

3) When running, the program generates files containing correction and

configuration data for Nexus. Files are named the Nexus serial number

plus extension frd and cfg. E.g. 1234567.frd.

4) If you are going to update more than  8-10 Nexus, the floppy disk

will run full. In that case it is necessary to copy the content of

the floppy disk to a seperate directory on your harddisk.

5) If you have some kind of breakdown during running the update don't be

too nervous, just restart the update program. The stored data will be

used.

6) If breakdown happens while extracting or storing data nothing will

be lost, because data in Nexus will not be reset before starting

downloading the program file. Just restart the update program.

7) It is not possible to update more than one Nexus at a time