Burster Resistomat Type 2318

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Burster Resistomat Type 2318.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Burster Resistomat Type 2318

instrument.  The functions appear in alphabetical order, with a description

of the function and its Visual Basic syntax, a description of each parameter,

and a list of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                    Function Name:

Initialize                        bp2318_init

Configure

Set Configure                  bp2318_set_config

Get Configure                  bp2318_get_config

Measure

Measure Start                  bp2318_meas_start

Measure Read                   bp2318_meas_read

Measure Stop                   bp2318_meas_stop

Comparator

Set Comparator Limit           bp2318_set_comp_limit

Get Comparator limit           bp2318_get_comp_limit

System

Set Address                    bp2318_set_address

Get Address                    bp2318_get_address

System Error                   bp2318_system_error

System Keyboard                bp2318_system_keyboard

Reset Device                   bp2318_reset

Read ID-String                 bp2318_read_id

Display Test                   bp2318_display_test

Send a SCPI-Message            bp2318_send_scpi

Receive a SCPI-Message         bp2318_receive_scpi

Close                             bp2318_close

________________________________________________________________________________

Burster Resistomat Type 2318

Instrument Name: Burster Resistomat Type 2318

Description:   This instrument module provides

RS232 support for Resistomat

Type 2318 from Burster.

Functions and Classes:

(1) Initialize :

Initialize the instrument and sets to

default configuration

(2) Configure:

Configures the instrument by setting the

range, correction, calibration, measurement

mode

(3) Measure

Starts measurement

Reads measurement value

Stops measurement

(4) Comparator:

Sets comparator limits and state

Gets comparator limits

(5) System:

Sets the group and user address

Gets the group and user address

Reads the system error

Resets the instrument

Reads the instrument ID-String

Controls the keyboard

Controls the display test

Sends a SCPI-Message

Receives a SCPI-Message

(6) Close :

Takes the instrument off-line.

*********

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

bp2318_close

Function bp2318_close& (ByVal instrumentID&)

Purpose

Panel Name :  Close an instrument

Description:  Closes a instrument on the RS232-

Port

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

221   Unable to close instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_display_test

Function bp2318_display_test& (ByVal instrumentID&)

Purpose

Panel Name:   Display Test

Description:  Activates or deactivates the

display test

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_get_address

Function bp2318_get_address& (ByVal instrumentID&, groupAddress&, userAddress&)

Purpose

Panel Name:    Get Address

Description:   Gets the group and user address

from the internal EEPROM

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

groupAddress&

Control Name:  Group Address

Description:   Displays the group address on the

RS232 (RS485) to be used to

communicate with the instrument.

Variable Type: Integer

userAddress&

Control Name:  Group Address

Description:   Displays the group address on the

RS232 (RS485) to be used to

communicate with the instrument.

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_get_comp_limit

Function bp2318_get_comp_limit& (ByVal instrumentID&, comparatorState&, lowerLimit#, upperLimit#)

Purpose

Panel Name:    Get Comparator Data

Description:   Gets the comparator limits and

comparator state

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

comparatorState&

Control Name:  Comparator state

Description:   Displays the comparator state.

Variable Type: Integer

Comparator state : 0 = OFF

1 = ON

lowerLimit#

Control Name:  Lower Limit

Description:   Displays the lower comparator

limit

Variable Type: Double

Unit : Ohm

upperLimit#

Control Name:  Upper Limit

Description:   Displays the upper comparator

limit

Variable Type: Double

Unit : Ohm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_get_config

Function bp2318_get_config& (ByVal instrumentID&, resistanceLoad&, autoZero&, range&, limit20mV&, tempComp&)

Purpose

Panel Name:    Get Configure

Description:   Initialize the instrument in

following ways:

- Gets the range of the instrument

- Gets the Auto-Zero state

- Gets the one of the three

possible temperature coeffic-

ients

- Gets the load to be measured

(mainly ohmic or mainly

inductive resistance)

- Gets the one of the three

possible 20mV limits.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

resistanceLoad&

Control Name:  Resistance Load

Description:   Displays the load to be measured.

Mainly ohmic or mainly inductive

resistance.

Variable Type: Integer

Resistance Load:   0 = Ohmic

1 = Inductive

autoZero&

Control Name:  Auto Zero

Description:   Displays the state of Auto-Zero

function.

Variable Type: Integer

Auto Zero:   0 = Auto Zero OFF

1 = Auto Zero ON

range&

Control Name:  Range

Description:   Displays the range of the

instrument .

Variable Type: Integer

Range:   0 = Auto Range

1 = 20  mOhm

2 = 200 mOhm

3 = 2   Ohm

4 = 20  Ohm

5 = 200 Ohm

6 = 2   KOhm

7 = 20  KOhm

8 = 200 KOhm

limit20mV&

Control Name:  20 mV Limit

Description:   Displays the range of the 20 mV

Limit.

Variable Type: Integer

20mV Limit:   0 = NONE

1 = 20 mV,1 A

2 = 20 mV,10 mA

3 = 20 mV,100 mA

tempComp&

Control Name:  Temp.Comp

Description:   Displays the temperature coeff-

itient of the temperature

compensation.

Variable Type: Integer

Temp.Compe:   0 = NONE

1 = CU/AL (3930 ppm/k)

2 = MS63  (1500 ppm/k)

3 = MS80  (1600 ppm/k)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

304   Unknown parameter from instrument

________________________________________________________________________________

bp2318_init

Function bp2318_init& (ByVal COMPort&, ByVal baudRate&, ByVal timeout_seconds#, ByVal blockCheck&, ByVal groupAddress&, ByVal userAddress&, ByVal resetDevice&, instrumentID&)

Purpose

Panel Name:    Initialize

Description:   Initializes the instrument in

following ways:

- opening a RS232-Port with user

specified baud-rate and time-out

- sets the instrument to a default

configuration

Parameter List

ByVal COMPort&

Control Name:   COM-Port

Description:    Specifies the COM-Port to be

used to communicate with the

instrument.

Variable Type:  Integer

Valid Values:   1 - COM1   (=> Adress 0x3f8 IRQ4)

2 - COM2   (=> Adress 0x2f8 IRQ3)

3 - COM3   (=> Adress 0x3e8 IRQ4)

4 - COM4   (=> Adress 0x2e8 IRQ3)

Default:        1

ByVal baudRate&

Control Name:   Baud Rate

Description:    Specifies the baud rate of the

RS232-Port to communicate with

the instrument. This value must

correspond to the current

internal setting of the

instrument.

Variable Type:  Integer

Valid Values:   1200,2400,4800,9600 or 19200 baud

Default:        9600

ByVal timeout_seconds#

Control Name:   Time-out (seconds)

Description:    Specifies  the time-out limit for

I/O operations.

A  value  of  0.0  will  disable

time-out altogether.

Variable Type:  Double

Valid Range:    0.060 to 100.0 seconds

Default:        1.0

ByVal blockCheck&

Control Name:   Bcc-Check

Description:    This parameter switches the block

check function of the driver ON

or OFF.

This value must correspond to the

current internal setting of the

instrument.

Variable Type:  Integer

Valid Values:   0 = OFF, 1 = ON

Default:        0 = OFF

ByVal groupAddress&

Control Name:  Group Address

Description:   Specifies the group address on the

RS232 (RS485) to be used to

communicate with the instrument.

This value must correspond to the

current internal group address

from the instrument.

Variable Type: Integer

Valid Range:   0 to 15

Default:       0

ByVal userAddress&

Control Name:  User Adress

Description:   Specifies the user address on the

RS232 (RS485) to be used to

communicate with the instrument.

This value must correspond to the

current internal user address

from the instrument.

Variable Type: Integer

Valid Range:   0 to 15

Default:       0

ByVal resetDevice&

Control Name:  Reset

Description:   Specifies if the instrument is to

be reset to its power-on settings

during the initialization

procedure.

Variable Type: Integer

Valid Range:   1 = Yes

0 = No

Default Value: 1 - Yes

NOTE: If you do not want the instrument reset

set this control to No while initializing

the instrument.

instrumentID&

Control Name:  Instrument ID

Description:   Returns an Instrument ID that is

used in all subsequent function

calls to select the instrument at

the RS232-Port and the address

specified.

If more than one instrument of the

same type is used, this value will

be used to differentiate between

them. The first instrument

initialized will be assigned an

Instrument ID of 1.

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   COM-Port is out of range

-2   Baud Rate is out of range

-3   Time-out is out of range

-4   Block Check parameter is out of range

-5   Group Address is out of range

-6   User Address is out of range

-7   Reset Device is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_meas_read

Function bp2318_meas_read& (ByVal instrumentID&, measuringValue#, ByVal unit$)

Purpose

Panel Name:    Measure Read

Description:   Reads the measuring values from

the instrument while a measurement

is done

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

measuringValue#

Control Name:  Measuring Value

Description:   Displays the measuring value.

The unit is Ohm

Variable Type: Double

ByVal unit$

Control Name:  Unit

Description:   Displays the measuring unit.

Variable Type: String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

302   Measuring isn't started

303   Overrange

________________________________________________________________________________

bp2318_meas_start

Function bp2318_meas_start& (ByVal instrumentID&)

Purpose

Panel Name:    Measure Start

Description:   Start the measurement

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_meas_stop

Function bp2318_meas_stop& (ByVal instrumentID&)

Purpose

Panel Name:    Measure Stop

Description:   Stops the measurement

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_read_id

Function bp2318_read_id& (ByVal instrumentID&, ByVal IDString$)

Purpose

Panel Name:   Read ID-String

Description:  Reads the ID-String from the

instrument

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal IDString$

Control Name:  ID-String

Description:   Displays the ID-String from the

instrument

Variable Type: String (Lenght = 35)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_receive_scpi

Function bp2318_receive_scpi& (ByVal instrumentID&, ByVal sendSCPIMessage$, ByVal instrumentAnswers$)

Purpose

Panel Name:   Receive a SCPI-Message

Description:  Sends "Query form" from SCPI-

Message and receive a answers

=> Polling

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal sendSCPIMessage$

Control Name:  Send SCPI-Message

Description:   This control is a string parameter

You must input Query-Form from

the SCPI-message without the

Line Feed termination.

Example      : ":SENS:CORR:STATE?"

Variable Type: String

ByVal instrumentAnswers$

Control Name:  Answer from Slave

Description:   This control return a string from

the slave.

Variable Type: String (Length => see the 2318

Manual SCPI-Commands)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_reset

Function bp2318_reset& (ByVal instrumentID&)

Purpose

Panel Name:   Reset device

Description:  Resets the instrument to the

default states

- Auto Range ON

- Load Ohmic

- 20mV Limit OFF

- Auto Zero OFF

- Temp. Compensation OFF

- Comparator OFF

- Local Lockout

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_send_scpi

Function bp2318_send_scpi& (ByVal instrumentID&, ByVal sendSCPIMessage$)

Purpose

Panel Name:   Send a SCPI-Message

Description:  Sends a SCPI-Message to a

slave and receive a ACK

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal sendSCPIMessage$

Control Name:  Send SCPI-Message

Description:   This control is a string parameter

You must input the SCPI-message

without the LineFeed termination.

Example      : ":SENS:CORR:STATE ON"

Variable Type: String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_set_address

Function bp2318_set_address& (ByVal instrumentID&, ByVal groupAddress&, ByVal userAddress&)

Purpose

Panel Name:    Set Address

Description:   Sets the group and user address in

the internal EEPROM

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal groupAddress&

Control Name:  Group Address

Description:   Specifies the group address on the

RS232 (RS485) to be used to

communicate with the instrument.

Variable Type: Integer

Valid Range:   0 to 15

Default:       0

ByVal userAddress&

Control Name:  User Address

Description:   Specifies the user address on the

RS232 (RS485) to be used to

communicate with the instrument.

Variable Type: Integer

Valid Range:   0 to 15

Default:       0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

-2   Group address is out of range

-3   User address is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_set_comp_limit

Function bp2318_set_comp_limit& (ByVal instrumentID&, ByVal comparatorState&, ByVal lowerLimit#, ByVal upperLimit#)

Purpose

Panel Name:    Comparator Limit

Description:   Set the comparator limits and the

comparator state

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal comparatorState&

Control Name:  Comparator State

Description:   This parameter activates or

deactivates the comparator

function.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default:       1 = ON

ByVal lowerLimit#

Control Name:  Lower Limit

Description:   Accepts the lower limit for the

comparator function. The unit is

"Ohm".

Variable Type: Double

Valid Range:   0.0 to  200000.00 (Ohm)

ByVal upperLimit#

Control Name:  Upper Limit

Description:   Accepts the upper limit for the

comparator function. The unit is

"Ohm".

Variable Type: Double

Valid Range:   0.0 to  200000.00 (Ohm)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

-2   Comparator state is out of range

-3   Lower limit is out of range

-4   Upper limit is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_set_config

Function bp2318_set_config& (ByVal instrumentID&, ByVal resistanceLoad&, ByVal autoZero&, ByVal range&, ByVal limit20mV&, ByVal tempComp&)

Purpose

Panel Name:    Set Configure

Description:   Initializes the instrument in

following ways:

- Sets the range of the instrument

- Switches the Auto-Zero function

ON/OFF

- Selects the one of the three

possible temperature coeffic-

ients

- Selects the load to be measured

(mainly ohmic or mainly

inductive resistance)

- Selects the one of the three

possible 20mV limits.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal resistanceLoad&

Control Name:  Resistance Load

Description:   Selects the load to be measured.

Mainly ohmic or mainly inductive

resistance.

Variable Type: Integer

Valid Range:   0 = Ohmic

1 = Inductive

ByVal autoZero&

Control Name:  Auto Zero

Description:   Switches the Auto Zero function

of the instrument ON/OFF.

Variable Type: Integer

Valid Range:   0 = Auto Zero OFF

1 = Auto Zero ON

ByVal range&

Control Name:  Range

Description:   Specifices the range of the

instrument .

Variable Type: Integer

Valid Range:   0 = Auto Range

1 = 20  mOhm

2 = 200 mOhm

3 = 2   Ohm

4 = 20  Ohm

5 = 200 Ohm

6 = 2   KOhm

7 = 20  KOhm

8 = 200 KOhm

ByVal limit20mV&

Control Name:  20 mV Limit

Description:   Selects the range of the 20 mV

Limit.

Variable Type: Integer

Valid Range:   0 = NONE

1 = 20 mV,1 A

2 = 20 mV,10 mA

3 = 20 mV,100 mA

ByVal tempComp&

Control Name:  Temp.Comp

Description:   Selects the temperature coeff-

itient for the temperature

compensation.

Variable Type: Integer

Valid Range:   0 = NONE

1 = CU/AL (3930 ppm/k)

2 = MS63  (1500 ppm/k)

3 = MS80  (1600 ppm/k)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

-2   Resistance load is out of range

-3   Auto Zero is out of range

-4   Range parameter is out of range

-5   20mV Limit parameter is out of range

-6   Temp.Comp. parameter is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_system_error

Function bp2318_system_error& (ByVal instrumentID&, ByVal systemError$)

Purpose

Panel Name:   System Error

Description:  Reads the system error of the

device

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal systemError$

Control Name:  System Error

Description:   Displays the SCPI system error

Variable Type: String

(see manual Resistomat Type 2318)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2318_system_keyboard

Function bp2318_system_keyboard& (ByVal instrumentID&, ByVal keyboard&)

Purpose

Panel Name:   System Keybroad

Description:  Controls the keybroad of the device

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  15

ByVal keyboard&

Control Name:  Keyboard

Description:   This parameter locks the keyboard

on the front panel of the

instrument

Variable Type: Integer

Valid Range:   0  = Unlocked

1  = Locked

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID is out of range

-2   Keyboard parameter is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

300   The instrument sends NAK

301   The instrument sends EOT