Burster Resistomat Type 2329

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Burster Resistomat Type 2329.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Burster Resistomat Type 2329

instrument.  The functions appear in alphabetical order, with a description

of the function and its C syntax, a description of each parameter, and a list

of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                         Function Name:

Initialize

Initialize GPIB                     bp2329_init_gpib

Initialize Serial                   bp2329_init_serial

Local (GPIB only)                   bp2329_local_gpib

Configure

Set Config Measurement              bp2329_set_config_meas

Get Config Measurement              bp2329_get_config_meas

Config Range

Set Config Range                 bp2329_set_config_range

Get Config Range                 bp2329_get_config_range

Set Config Auto Range            bp2329_set_config_auto_range

Get Config Auto Range            bp2329_get_config_auto_range

Config Display

Set Config Display Mode          bp2329_set_config_display_mode

Get Config Display Mode          bp2329_get_config_display_mode

Set Config Display Options       bp2329_set_config_display_options

Get Config Display Options       bp2329_get_config_display_options

Set Config Display Contrast      bp2329_set_config_display_contrast

Get Config Display Contrast      bp2329_get_config_display_contrast

Config Temperature Compensation

Set Config Temp. Compensation    bp2329_set_config_temp_compensation

Get Config Temp. Compensation    bp2329_get_config_temp_compensation

Get Temperature Value            bp2329_get_temp_value

Set Config Scale  PT100          bp2329_set_config_scale_pt100

Get Config Scale  PT100          bp2329_get_config_scale_pt100

Set Config Scale  U/I Input      bp2329_set_config_scale_ui_input

Get Config Scale  U/I Input      bp2329_get_config_scale_ui_input

Set Config Temp. Coefficient     bp2329_set_config_temp_coefficient

Get Config Temp. Coefficient     bp2329_get_config_temp_coefficient

Set Active Temp. Coefficient     bp2329_set_active_temp_coefficient

Get Active Temp. Coefficient     bp2329_get_active_temp_coefficient

Set Config Device Access            bp2329_set_config_device_access

Get Config Device Access            bp2329_get_config_device_access

Config Printer

Set Config Printer State         bp2329_set_config_printer_state

Set Config Printer Mode          bp2329_set_config_printer_mode

Get Config Printer Mode          bp2329_get_config_printer_mode

Reset Printer Counter Reset      bp2329_printer_counter_reset

Store/Load Configuration

Set Memory Label                 bp2329_set_config_memory_label

Get Memory Label                 bp2329_get_config_memory_label

Set Memory No.                   bp2329_set_config_memory

Get Memory No.                   bp2329_get_config_memory

Measure

Measure Start                       bp2329_meas_start

Measure Read                        bp2329_meas_read

Measure Read String                 bp2329_meas_read_string

Measure Stop                        bp2329_meas_stop

Comparator

Set Config Comparator Mode          bp2329_set_config_comparator_mode

Get Config Comparator Mode          bp2329_get_config_comparator_mode

Set Config 2 Limit Comparator       bp2329_set_config_2_limit_comparator

Get Config 2 Limit Comparator       bp2329_get_config_2_limit_comparator

Set Config 4 Limit Comparator       bp2329_set_config_4_limit_comparator

Get Config 4 Limit Comparator       bp2329_get_config_4_limit_comparator

Read 2 Limit Comparator Report      bp2329_comparator_2_limit_report

Read 4 Limit Comparator Report      bp2329_comparator_4_limit_report

Clear Comparator                    bp2329_comparator_clear

Envelope (Min/Max-Function)

Set Envelope State                  bp2329_set_envelope_state

Get Envelope State                  bp2329_get_envelope_state

Read Envelope Report                bp2329_envelope_read_report

Clear Envelope                      bp2329_envelope_clear

Datalogger

Set Datalogger State                bp2329_set_datalogger_state

Get Datalogger State                bp2329_get_datalogger_state

Set Datalogger Name                 bp2329_set_datalogger_name

Get Datalogger Name                 bp2329_get_datalogger_name

Set Datalogger Block Size           bp2329_set_datalogger_block_size

Get Datalogger Block Size           bp2329_get_datalogger_block_size

Set Datalogger Filter               bp2329_set_datalogger_filter

Get Datalogger Filer                bp2329_get_datalogger_filter

Clear Datalogger Block              bp2329_datalogger_clear_block

Set Active Datalogger Block         bp2329_set_active_datalogger_block

Get Active Datalogger Block         bp2329_get_active_datalogger_block

Read Datalogger Num. of Counts      bp2329_datalogger_read_num_counts

Read Num. of Datalogger Block       bp2329_datalogger_read_block

Read Datalogger Eval. Values        bp2329_datalogger_read_eval_values

System

Set Register                        bp2329_system_set_register

Get Register                        bp2329_system_get_register

System Date/Time                    bp2329_system_date_time

System Error                        bp2329_system_error

System Keyboard                     bp2329_system_keyboard

Calibration

Calibration State                   bp2329_cal_state

Calibration Resistance Ranges       bp2329_cal_fres

Calibration 20mV Mode               bp2329_cal_20mV_mode

Calibration 20mV Init               bp2329_cal_20mV_init

Calibration 20mV                    bp2329_cal_20mV

Calibration Voltage Mode            bp2329_cal_voltage_mode

Calibration Voltage                 bp2329_cal_voltage

Calibration Current                 bp2329_cal_current

Calibration PT100 Mode              bp2329_cal_pt100_mode

Calibration PT100                   bp2329_cal_pt100

SCPI

Send a SCPI-Message                 bp2329_send_scpi

Receive a SCPI-Message              bp2329_receive_scpi

Close                                  bp2329_close

________________________________________________________________________________

Burster Resistomat Type 2329

Instrument Name:Burster Resistomat Typ 2329

Description:   This instrument module provides

RS232 support for  Resistomat

Typ 2329 from Burster.

The module contains the following

functions:

(1) Initialize Functions

Initialize the GPIB

Initialize the RS232-Port

(2) Configure  the device   - Range

- Display

- Temp.Compensation

- Device Access

- Printer

- Store/Load Config.

(3) Measurement Functions

Start Measurement

Read Measurement value

Stop Measurement

(4) Comparartor Functions

Configure Comparator Mode

Configure 2 Limit Mode

Configure 4 Limit Mode

Read Comparator Reports

Clear Comparator

(5) Envelope (Min/Max-Functions)

Envelope State

Read Envalope Report

Clear Enevlope

(6) Datalogger

Configure Datalogger

Datalogger State

Clear Datalogger Block

Select Active Datalogger Block

Read 20 Datalogger Values

Read Num. of Datalogger Block

Read Datalogger Evaluation Values

(7) System Functions

Set Register

Get Register

Set Device System Date and Time

Read the system error

Control the keyboard

(8) Calibration Functions

(9) Close the instrument

*********

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

bp2329_cal_20mV

int bp2329_cal_20mV (int instrumentID,

double resistanceValue_ca200_mOhm,

double lowVoltage, double highVoltage);

Purpose

Panel Name:   Calibration 20mV

Descrpition:  This command enter the resistance value and the low

voltage

and the high voltage to the device.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

resistanceValue_ca200_mOhm

Variable Type       double

Control Name:  Resistance Value (ca. 200mOhm)

Description:   Input the resistance value

Variable Type: double

lowVoltage

Variable Type       double

Control Name:  Low Voltage

Description:   Input the low voltage value from

the cal_20mV_init measure before.

Variable Type: double

highVoltage

Variable Type       double

Control Name:  High Voltage

Description:   Input the high voltage value from

the cal_20mV_init measure before.

Variable Type: double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_20mV_init

int bp2329_cal_20mV_init (int instrumentID);

Purpose

Panel Name:   Calibration 20mV Init

Descrpition:  Initialize the 20mV calibration.

Before you must select the 20mV mode.

Please measure the voltage about the

calibration resistance during both

20mV modes. After this command enter

the resistance value and the low voltage

and the high voltage to the device

with following function.

bp2329_cal_20mV(id,res_value,low_volatge,

high_voltage);

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_20mV_mode

int bp2329_cal_20mV_mode (int instrumentID, int mode);

Purpose

Panel Name:   Calibration 20mV Mode

Descrpition:  Specificies the 20mV Mode (High or Low)

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

mode

Variable Type       int

Control Name:  Mode

Description:   Select the 20mV calibration mode

Variable Type: Integer

Valid Range:   0  = LOW

1  = High

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid mode

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_current

int bp2329_cal_current (int instrumentID, double currentValue);

Purpose

Panel Name:   Calibration Current

Descrpition:  Enter the current value (20mA)

to finished the calibration of current.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

currentValue

Variable Type       double

Control Name:  Current Value

Description:   Input the current value to calibrate.

Variable Type: double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_fres

int bp2329_cal_fres (int instrumentID, int range, double value);

Purpose

Panel Name:   Calibration Resistance Range

Descrpition:  This panel calibrate all possible

resistance ranges

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

range

Variable Type       int

Control Name:  Range

Description:   Select the range to calibrate

Variable Type: Integer

Valid Range:   0  = 200mOhm

1  = 2Ohm

2  = 20Ohm

3  = 200Ohm

4  = 2KOhm

5  = 20KOhm

6  = 200KOhm

value

Variable Type       double

Control Name:  Value

Description:   Input the calibration value for

the selected range

Variable Type: double

Valid Range:   200mOhm - 200 KOHm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid range

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_pt100

int bp2329_cal_pt100 (int instrumentID);

Purpose

Panel Name:   Calibration PT100

Descrpition:  Initialize the PT100 calibration

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_pt100_mode

int bp2329_cal_pt100_mode (int instrumentID, int mode, double value);

Purpose

Panel Name:   Calibration PT100 Mode

Descrpition:  Calibrate the PT100 Input of the device.

Enter the lwo and high value to

calibrate the PT100 input.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

mode

Variable Type       int

Control Name:  Mode

Description:   Select the PT100 calibration mode

Variable Type: Integer

Valid Range:   0  = LOW

1  = High

value

Variable Type       double

Control Name:  Value

Description:   Input the PT100 calibration value.

Variable Type: Integer

Valid Range:   LOW-Mode = ca 50 Ohm

High-Mode = ca. 200 Ohm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid mode

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_state

int bp2329_cal_state (int instrumentID, int state);

Purpose

Panel Name:   Calibration State

Descrpition:  Sets the calibration state ON/OFF.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

state

Variable Type       int

Control Name:  State

Description:   If the State is On the calibration

mode is active

Variable Type: Integer

Valid Range:   0  = OFF

1  = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid State

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_voltage

int bp2329_cal_voltage (int instrumentID, double lowVoltage,

double highVoltage);

Purpose

Panel Name:   Calibration Voltage

Descrpition:  Enter the low voltage (0V) and the

high voltage (10V) to finished the

calibration of voltage.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

lowVoltage

Variable Type       double

Control Name:  Low Voltage

Description:   Input the low voltage value from

the cal_volatge_mode measure before.

Variable Type: double

highVoltage

Variable Type       double

Control Name:  High Voltage

Description:   Input the high voltage value from

the cal_volatge_mode measure before.

Variable Type: double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_cal_voltage_mode

int bp2329_cal_voltage_mode (int instrumentID, int mode);

Purpose

Panel Name:   Calibration Voltage Mode

Descrpition:  Initialize the voltage calibration.

Input 0V and sets the mode = low.

Input 10V and sets the mode = high.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

mode

Variable Type       int

Control Name:  Mode

Description:   Select the voltage calibration mode

Variable Type: Integer

Valid Range:   0  = LOW

1  = High

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Mode

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_close

int bp2329_close (int instrumentID);

Purpose

Panel Name :  Close

Descrpition:  Close the instrument

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_comparator_2_limit_report

int bp2329_comparator_2_limit_report (int instrumentID, int *value1,

int *value2, int *value3);

Purpose

Panel Name:    Read 2 Limit Comparator Report

Description:   Read the follow the comparator

values as integer

- Value 1 (Num.of values < lower limit)

- Value 2 (lower limit >= Num.of values <= upper limit)

- Value 3 (Num.of values > upper limit)

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

value1

Variable Type       int (passed by reference)

Control Name:  Value 1

Description:   Value 1 returns the Numbers of values

are smaller than lower limit.

Variable Type: Integer

value2

Variable Type       int (passed by reference)

Control Name:  Value 2

Description:   Value 2 returns the Numbers of values

are greather than lower limit and smaller

than upper limit.

Variable Type: Integer

value3

Variable Type       int (passed by reference)

Control Name:  Value 3

Description:   Value 3 returns the Numbers of values

are greather than upper limit.

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

310   Device Voltage Error

311   Device Current Error

312   Device Temperature Error

313   Device Calibration Error

314   Device In Over Error

315   Device Out Over Error

316   Device Lowbat Error

________________________________________________________________________________

bp2329_comparator_4_limit_report

int bp2329_comparator_4_limit_report (int instrumentID, int *value1,

int *value2, int *value3,

int *value4, int *value5);

Purpose

Panel Name:    Read 4 Limit Comparator Report

Description:   Read the follow the comparator

values as integer

- Value 1 (Num.of values < limit1)

- Value 2 (limit1>= Num.of values < limit2)

- Value 3 (limit2>= Num.of values < limit3)

- Value 4 (limit3>= Num.of values < limit4)

- Value 5 (Num.of values > limit4)

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

value1

Variable Type       int (passed by reference)

Control Name:  Value 1

Description:   Value 1 returns the Numbers of values

are smaller than limit1.

Variable Type: Integer

value2

Variable Type       int (passed by reference)

Control Name:  Value 2

Description:   Value 2 returns the Numbers of values

are greather than limit1 and smaller

than limit2.

Variable Type: Integer

value3

Variable Type       int (passed by reference)

Control Name:  Value 3

Description:   Value 3 returns the Numbers of values

are greather than limit3 and smaller

than limit3.

Variable Type: Integer

value4

Variable Type       int (passed by reference)

Control Name:  Value 4

Description:   Value 4 returns the Numbers of values

are greather than limit3 and smaller

than limit4.

Variable Type: Integer

value5

Variable Type       int (passed by reference)

Control Name:  Value 5

Description:   Value 5 returns the Numbers of values

are greather than limit4.

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

310   Device Voltage Error

311   Device Current Error

312   Device Temperature Error

313   Device Calibration Error

314   Device In Over Error

315   Device Out Over Error

316   Device Lowbat Error

________________________________________________________________________________

bp2329_comparator_clear

int bp2329_comparator_clear (int instrumentID);

Purpose

Panel Name:    Clear Comparator

Description:   Clear the comparator datas

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_datalogger_clear_block

int bp2329_datalogger_clear_block (int instrumentID, int block);

Purpose

Panel Name:    Clear Datalogger Block

Description:   Clear the datalogger block no.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

block

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to clear

Variable Type: Integer

Valid Range:   0 - 31

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_datalogger_read_block

int bp2329_datalogger_read_block (int instrumentID, int blockNo,

int startNumber, char value[],

char date[], char time[]);

Purpose

Panel Name:    Read Datalogger Block

Description:   Read measure,date and time values

from the specificies block no.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to read

Variable Type: Integer

Valid Range:   0 - 31

startNumber

Variable Type       int

Control Name:  Start Number

Description:   Specificies the start number from to you

want to read 20 Values

Variable Type: Integer

Valid Range:   1 - 20000

value

Variable Type       char []

Control Name:  Values

Description:   return the meas.values of the block.

Variable Type: String

date

Variable Type       char []

Control Name:  Date

Description:   return the date of the block.

Variable Type: String

time

Variable Type       char []

Control Name:  Time

Description:   return the time of the block.

Variable Type: String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_datalogger_read_eval_values

int bp2329_datalogger_read_eval_values (int instrumentID, int blockNo,

char minimum[], char maximum[],

char meanValue[],

char stdDeviation[]);

Purpose

Panel Name:    Read Datalogger Eval.Block

Description:   Read evaluation values (min,max,mean,std-dev.,

from the specificies block no.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to read

Variable Type: Integer

Valid Range:   0 - 31

minimum

Variable Type       char []

Control Name:  Minimum

Description:   Return the minimum value of the block.

Variable Type: String

maximum

Variable Type       char []

Control Name:  Maximum

Description:   Return the maximum value of the block.

Variable Type: String

meanValue

Variable Type       char []

Control Name:  Mean value

Description:   Return the mean value of the block.

Variable Type: String

stdDeviation

Variable Type       char []

Control Name:  Std. Deviation

Description:   Return the Std.deviation of the block.

Variable Type: String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_datalogger_read_num_counts

int bp2329_datalogger_read_num_counts (int instrument_ID, int blockNo,

int *count);

Purpose

Panel Name:    Read Datalogger Num.Of Counts

Description:   Read the num. of counts from the block.

Parameter List

instrument_ID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to clear

Variable Type: Integer

Valid Range:   0 - 31

count

Variable Type       int (passed by reference)

Control Name:  Count

Description:   Returns the numbers of counts from

the block.

Variable Type: Integer

Valid Range:   0 = Off

1 = On

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_envelope_clear

int bp2329_envelope_clear (int instrumentID);

Purpose

Panel Name:    Clear Envelope

Description:   Clear the envelope datas

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_envelope_read_report

int bp2329_envelope_read_report (int instrumentID, char minValue[],

char maxValue[], char diffValue[]);

Purpose

Panel Name:    Read Envelope Report

Description:   Read the follow the envelope

values as String

- Minium Value

- Maximum Value

- Difference Value

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

minValue

Variable Type       char []

Control Name:  Min. Value

Description:   Returns the minimum envelope

Value of the instrument.

Variable Type: string

maxValue

Variable Type       char []

Control Name:  Max. Value

Description:   Returns the maximum envelope

Value of the instrument.

Variable Type: string

diffValue

Variable Type       char []

Control Name:  Diff. Value

Description:   Returns the difference envelope

Value of the instrument.

Variable Type: string

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

310   Device Voltage Error

311   Device Current Error

312   Device Temperature Error

313   Device Calibration Error

314   Device In Over Error

315   Device Out Over Error

316   Device Lowbat Error

________________________________________________________________________________

bp2329_get_active_datalogger_block

int bp2329_get_active_datalogger_block (int instrumentID, int *blockNo);

Purpose

Panel Name:    Get Active Datalogger Block

Description:   Gets Active Datalogger Block

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int (passed by reference)

Control Name:  Block No.

Description:   Returns the active block no.

Variable Type: Integer

Valid Range:   0 - 31

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_active_temp_coefficient

int bp2329_get_active_temp_coefficient (int instrumentID,

int *coeffValue);

Purpose

Panel Name:    Get Active Temperature Coefficients

Descrpition:   Get the active temperature

coefficents.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

coeffValue

Variable Type       int (passed by reference)

Control Name:  Value of coefficient

Description:   This parameter returns the acvtive

temperature coefficients

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Number of TC

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_2_limit_comparator

int bp2329_get_config_2_limit_comparator (int instrumentID,

double *lowerLimit,

double *upperLimit);

Purpose

Panel Name:    Get Two Limit Comparator

Description:   Gets the two limit comparator

parameter

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

lowerLimit

Variable Type       double (passed by reference)

Control Name:  Lower Limit

Description:   Gets the lower limit of the

two limit comparator

Variable Type: Double

Default:       0.0

upperLimit

Variable Type       double (passed by reference)

Control Name:  Upper Limit

Description:   Gets Specifies the upper limit of the

two limit comparator

Variable Type: Double

Default:       0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_get_config_4_limit_comparator

int bp2329_get_config_4_limit_comparator (int instrumentID,

double *limit1,

double *limit2,

double *limit3,

double *limit4);

Purpose

Panel Name:    Get Four Limit Comparator

Description:   Gets the four limit

comparator parameter

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

limit1

Variable Type       double (passed by reference)

Control Name:  Limit 1

Description:   Gets the limit 1 of the

four limit comparator

Variable Type: Double

Default:       0.0

limit2

Variable Type       double (passed by reference)

Control Name:  Limit2

Description:   Gets the limit 2 of the

four limit comparator

Variable Type: Double

Default:       0.0

limit3

Variable Type       double (passed by reference)

Control Name:  Limit 3

Description:   Gets the limit 3 of the

four limit comparator

Variable Type: Double

Default:       0.0

limit4

Variable Type       double (passed by reference)

Control Name:  Limit4

Description:   Gets the limit 4 of the

four limit comparator

Variable Type: Double

Default:       0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_get_config_auto_range

int bp2329_get_config_auto_range (int instrumentID, int *lowerAutoRange,

int *upperAutoRange);

Purpose

Panel Name:    Gets Config Auto Range

Description:   Gets the upper and lower range for

the auto range function of the device

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

lowerAutoRange

Variable Type       int (passed by reference)

Control Name:  Lower Auto Range

Description:   This parameter gets the lowest

range in automatic measurement

range selection

Variable Type: Integer

Valid Range:   1 = 200 mOhm

2 = 2    Ohm

3 = 20   Ohm

4 = 200  Ohm

5 = 2   KOhm

6 = 20  KOhm

7 = 200 KOhm

upperAutoRange

Variable Type       int (passed by reference)

Control Name:  Upper Auto Range

Description:   This parameter gets the highest

range in automatic measurement

range selection

Variable Type: Integer

Valid Range:   1 = 200 mOhm

2 = 2    Ohm

3 = 20   Ohm

4 = 200  Ohm

5 = 2   KOhm

6 = 20  KOhm

7 = 200 KOhm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid lower Range

-3   Invalid upper range

-4   upper range <= lower range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_comparator_mode

int bp2329_get_config_comparator_mode (int instrumentID, int *state,

int *relaisMode, int *onError);

Purpose

Panel Name:    Get Comparator Mode

Description:   Gets the Comparator State ON/OFF

Config the relais mode

Config the comparator fault mode

(On Error)

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

state

Variable Type       int (passed by reference)

Control Name:  State

Description:   Gets the state is On the comparator

Variable Type: Integer

Valid Range:   0 = Off

1 = On

relaisMode

Variable Type       int (passed by reference)

Control Name:  Relais Mode

Description:   Gets the relais mode

Variable Type: Integer

Valid Range:   0 = Off

1 = On

onError

Variable Type       int (passed by reference)

Control Name:  On Error

Description:   Gets the On Error is upper,the

comparator show the ">" - token

by an error

Variable Type: Integer

Valid Range:   0 = None

1 = Upper ">"-token

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

-3   Invalid relais mode

-4   Invalid on error

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_device_access

int bp2329_get_config_device_access (int instrumentID, int *measureMenu,

int *displayMenu, int *printerMenu,

int *tempCompMenu,

int *contrastMenu,

int *storeConfigMenu,

int *interfaceMenu,

int *displayStateMenu,

int *auxInOutMenu, int *scaleMenu,

int *dataloggerMenu,

int *calibrationMenu,

int *dateMenu, int *fastStartMenu,

int *comparaConfMenu,

int *comparaEvalMenu,

int *envelopeConfMenu,

int *envelopeEvalMenu,

int *manRangeMenu);

Purpose

Panel Name:    Config Device Access

Description:   Configure the access to the device

menues.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

measureMenu

Variable Type       int (passed by reference)

Control Name:  Measure Menu

Description:   This parameter gets the state of

measure menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

displayMenu

Variable Type       int (passed by reference)

Control Name:  Display Menu

Description:   This parameter gets the state of the

display menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

printerMenu

Variable Type       int (passed by reference)

Control Name:  Printer Menu

Description:   This parameter gets the state of the

Printer menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

tempCompMenu

Variable Type       int (passed by reference)

Control Name:  Temp. Comp. Menu

Description:   This parameter gets the state of the

temperature compensation menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

contrastMenu

Variable Type       int (passed by reference)

Control Name:  Contrast Menu

Description:   This parameter gets the state of the

contrast menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

storeConfigMenu

Variable Type       int (passed by reference)

Control Name:  Store Config. Menu

Description:   This parameter gets the state of the

store config. menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

interfaceMenu

Variable Type       int (passed by reference)

Control Name:  Interface Menu

Description:   This parameter gets the state of the

interface menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

displayStateMenu

Variable Type       int (passed by reference)

Control Name:  Display State Menu

Description:   This parameter gets the state of the

display state menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

auxInOutMenu

Variable Type       int (passed by reference)

Control Name:  Aux. In/Out Menu

Description:   This parameter gets the state of the

Aux. Input/Output menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

scaleMenu

Variable Type       int (passed by reference)

Control Name:  Scale Menu

Description:   This parameter gets the state of the

scale menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

dataloggerMenu

Variable Type       int (passed by reference)

Control Name:  Datalogger Menu

Description:   This parameter gets the state of the

datalogger menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

calibrationMenu

Variable Type       int (passed by reference)

Control Name:  Calibration Menu

Description:   This parameter gets the state of the

Calibration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

dateMenu

Variable Type       int (passed by reference)

Control Name:  Date Menu

Description:   This parameter gets the state of the

date menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

fastStartMenu

Variable Type       int (passed by reference)

Control Name:  Fast Start Menu

Description:   This parameter gets the state of the

fast start menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

comparaConfMenu

Variable Type       int (passed by reference)

Control Name:  Comparator Config Menu

Description:   This parameter gets the state of the

comparator config menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

comparaEvalMenu

Variable Type       int (passed by reference)

Control Name:  Comparator Eval. Menu

Description:   This parameter gets the state of the

comparator eval. menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

envelopeConfMenu

Variable Type       int (passed by reference)

Control Name:  Envelope Config. Menu

Description:   This parameter gets the state of the

envelope config (Envelope Start) menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

envelopeEvalMenu

Variable Type       int (passed by reference)

Control Name:  Envelope Eval. Menu

Description:   This parameter gets the state of the

envelope evalation menu(Reset Button).

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

manRangeMenu

Variable Type       int (passed by reference)

Control Name:  Manual Range Menu

Description:   This parameter gets the state of the

manual range menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid meaure menu

-3     Invalid display menu

-4     Invalid printer menu

-5     Invalid temp.comp menu

-6     Invalid contrast menu

-7     Invalid store config. menu

-8     Invalid interface menu

-9     Invalid display state menu

-10     Invalid aux. input/output menu

-11     Invalid scale menu

-12     Invalid datalogger menu

-13     Invalid calibration menu

-14     Invalid date menu

-15     Invalid fast start menu

-16     Invalid comparator config. menu

-17     Invalid comparator eval. menu

-18     Invalid envelope config. menu

-19     Invalid envelope eval. menu

-20     Invalid manuall range menu

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_display_contrast

int bp2329_get_config_display_contrast (int instrumentID,

double *displayContrast);

Purpose

Panel Name:    Get Config Display Contrast

Descrpition:   Gets the display contrast

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

displayContrast

Variable Type       double (passed by reference)

Control Name:  Display Contrast

Description:   This parameter gets the

adjusting of the display contrast.

Variable Type: Double

Valid Range:   0.0 - 1.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Contrast Value

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_display_mode

int bp2329_get_config_display_mode (int instrumentID, int *displayMode,

double *deltaNominalValue,

int *resolution);

Purpose

Panel Name:    Get Config Display Mode

Description:

- Gets the display mode

- Gets the display resolution

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

displayMode

Variable Type       int (passed by reference)

Control Name:  Display Mode

Description:   This parameter select the display mode.

Variable Type: Integer

Valid Range:   0 = Ohm

1 = Delta %

2 = Comparator

deltaNominalValue

Variable Type       double (passed by reference)

Control Name:  Display Nominal Value

Description:   This parameter inputs the nominal value

from the delta% display mode.

Variable Type: double

resolution

Variable Type       int (passed by reference)

Control Name:  Resolution

Description:   This parameter select the resolution

of the device.

Variable Type: Integer

Valid Range:   0 = 2000 Digits

1 = 20000 Digits

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid display mode

-3     Invalid resolution

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_display_options

int bp2329_get_config_display_options (int instrumentID,

int *rangeOption,

int *currentOption,

int *temperatureOption,

int *sampleOption,

int *triggerOption,

int *autoManOption,

int *loadOption,

int *measureOption);

Purpose

Panel Name:    Get Config Display Options

Descrpition:   Get the display options of the

device

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

rangeOption

Variable Type       int (passed by reference)

Control Name:  Range Option

Description:   This parameter gets the state

range function of the

device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

currentOption

Variable Type       int (passed by reference)

Control Name:  Current Option

Description:   This parameter gets the state

of the current function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

temperatureOption

Variable Type       int (passed by reference)

Control Name:  Temperature Option

Description:   This parameter gets

the temperature function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

sampleOption

Variable Type       int (passed by reference)

Control Name:  Sample Option

Description:   This parameter gets the state

of the sample function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

triggerOption

Variable Type       int (passed by reference)

Control Name:  Trigger Option

Description:   This parameter gets the state

of the trigger function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

autoManOption

Variable Type       int (passed by reference)

Control Name:  Auto/Man Option

Description:   This parameter gets the state

of the Auto/Man function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

loadOption

Variable Type       int (passed by reference)

Control Name:  Load Option

Description:   This parameter gets the state

of the load function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

measureOption

Variable Type       int (passed by reference)

Control Name:  Measure Option

Description:   This parameter gets the state

of the measure function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid range option

-3     Invalid current option

-4     Invalid temperature option

-5     Invalid sample option

-6     Invalid trigger option

-7     Invalid auto/man option

-8     Invalid load option

-9     Invalid measure option

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_meas

int bp2329_get_config_meas (int instrumentID, int *averageMode,

int *averageCounts, int *loading,

int *measurementMode, int *_20mVLimit,

int *measurementSequence,

int *num_ofConvertsCycles);

Purpose

Panel Name:    Get Config Device

Description:

- Gets the average counts

- Gets the average control

- Gets the loading mode

- Gets the measurement mode

- Gets the 20 mV limit

- Gets the measurement sequence

- Gets the numbers of convert Cycles

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

averageMode

Variable Type       int (passed by reference)

Control Name:  Average Control

Description:   This parameter gets the type of

Averaging

Variable Type: Integer

Valid Range:   0 = Moving

1 = Repeat

averageCounts

Variable Type       int (passed by reference)

Control Name:  Average Counts

Description:   This parameter gets the numbers

of average counts

Variable Type: Integer

Valid Range:   1 - 9999

loading

Variable Type       int (passed by reference)

Control Name: Loading

Description: This parameter gets the load to

measured (mainly ohmic or mainly

indictive resistance).

Variable Type: Integer

Valid Range:   0 = mainly ohmic

1 = mainly inductive

measurementMode

Variable Type       int (passed by reference)

Control Name:  Measurement Mode

Description:   This parameter gets the

measurements mode.

Variable Type: Integer

Valid Range:   0 = Single

1 = Repetition

_20mVLimit

Variable Type       int (passed by reference)

Control Name:  20 mV Limit

Description:   Gets the range of the 20 mV

Limit.

Variable Type: Integer

Valid Range:   0 = Off

1 = ON

measurementSequence

Variable Type       int (passed by reference)

Control Name:  Measurement Sequence

Description:   This parameter gets then measurement

sequence.

Variable Type: Integer

Valid Range:   0 = Ref-Comp

(Zero measuremnet with current = 0)

1 = None Comp

(Measurement without zero measurement)

2 = Single Comp

(Measurement with one zero measurement)

3 = Standard

(Measurement with toggle between zero

measurement and realy measurement)

num_ofConvertsCycles

Variable Type       int (passed by reference)

Control Name:  Number of Converts

Description:   This parameter gets the number of converts.

Variable Type: Integer

Valid Range:   0 = Minimum

1 = Medium

2 = Maximum

3 = Standard

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Average Control

-3   Invalid Average Counts

-4   Invalid Loading Mode

-5   Invalid Measurement Mode

-6   Invalid 20 mV State

-7   Invalid Measurement Sequence

-8   Invalid Numbers of Convert Cycles

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_memory

int bp2329_get_config_memory (int instrumentID, int memoryNo);

Purpose

Panel Name:    Get Memory No.

Description:   Recall the actual Parameters to

the selected memory of the

instrument

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

memoryNo

Variable Type       int

Control Name:   Memory No.

Description:    Specifies the Memory No.

Variable Type:  integer

Valid Range:    0 to 32

Default:        0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Memory No.

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

320   Datalogger is empty

________________________________________________________________________________

bp2329_get_config_memory_label

int bp2329_get_config_memory_label (int instrumentID, int memoryNo,

char memoryLabel[]);

Purpose

Panel Name:    Get Memory Label

Description:   Return the memory label from the

specificies memory No.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

memoryNo

Variable Type       int

Control Name:  Memory No.State

Description:   Specificies the memory No. to

store the the label text

Variable Type: Integer

Valid Range:   0 -  32

memoryLabel

Variable Type       char []

Control Name:  Memory Label

Description:   Return the memory label.

Variable Type: String (10 Byte + NULL)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_printer_mode

int bp2329_get_config_printer_mode (int instrumentID,

int *printTemperature,

int *printCounter, int *printLabel,

int *printTime, int *printDate,

int *measValuesHeader,

char labelText[], int *HH, int *MM,

int *SS);

Purpose

Panel Name:    Set Config Printer Mode

Description:   Configure the printer options

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

printTemperature

Variable Type       int (passed by reference)

Control Name:  Print Temperature

Description:   Switch the print temperature function

ON or OFF.

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printCounter

Variable Type       int (passed by reference)

Control Name:  Print Counter

Description:   Switch the print counter function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printLabel

Variable Type       int (passed by reference)

Control Name:  Print Label

Description:   Switch the print label function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printTime

Variable Type       int (passed by reference)

Control Name:  Print Time

Description:   Switch the print time function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printDate

Variable Type       int (passed by reference)

Control Name:  Print Date

Description:   Switch the print date function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

measValuesHeader

Variable Type       int (passed by reference)

Control Name:  Meas.Values/Header

Description:   Specifices how many measure values

do want print per header.

Variable Type: Integer

Valid Range:   1 - 1000

labelText

Variable Type       char []

Control Name:  Label text

Description:   Specifices the label text for the

printer.

Variable Type: String

Valid Range:   max. 10 Byte

HH

Variable Type       int (passed by reference)

Control Name:  Timer Interval HH

Description:   Specifices the hours for the timer

interval.

Variable Type: Integer

Valid Range:   0 -99

MM

Variable Type       int (passed by reference)

Control Name:  Timer Interval MM

Description:   Specifices the minutes for the timer

interval.

Variable Type: Integer

Valid Range:   0 -59

SS

Variable Type       int (passed by reference)

Control Name:  Timer Interval SS

Description:   Specifices the seconds for the timer

interval.

Variable Type: Integer

Valid Range:   0 - 59

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Print Temperature

-3   Invalid Print Counter

-4   Invalid Print Label

-5   Invalid Print Time

-6   Invalid Print Date

-7   Invalid Meas.Values/Header

-8   Invalid Label Text (to long)

-9   Invalid Timer Interval HH

-10   Invalid Timer Interval MM

-11   Invalid Timer Interval SS

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_range

int bp2329_get_config_range (int instrumentID, int *range);

Purpose

Panel Name:    Get Config Range

Description:   Gets the range

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

range

Variable Type       int (passed by reference)

Control Name:  Range

Description:   This parameter gets measurement range

Variable Type: Integer

Valid Range:

1 = 200 mOhm

2 = 2    Ohm

3 = 20   Ohm

4 = 200  Ohm

5 = 2   KOhm

6 = 20  KOhm

7 = 200 KOhm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_scale_pt100

int bp2329_get_config_scale_pt100 (int instrumentID, double *r0,

double *coefficienta,

double *coefficientb);

Purpose

Panel Name:    Get Config Scale PT100

Description:   Gets the scale parameter of the

PT100 function

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

r0

Variable Type       double (passed by reference)

Control Name:  R0

Description:   Gets R0

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Variable Type: Double

Default:       0.0

coefficienta

Variable Type       double (passed by reference)

Control Name:  Coefficient "a"

Description:   Gets the coefficient "a".

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Variable Type: Double

Default:       0.0

coefficientb

Variable Type       double (passed by reference)

Control Name:  Coefficient "b"

Description:   Gets the coefficient "b".

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Variable Type: Double

Default:       0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_get_config_scale_ui_input

int bp2329_get_config_scale_ui_input (int instrumentID, int mode,

double *u1I1, double *u2I2,

double *t1, double *t2);

Purpose

Panel Name:    Config Scale U/I Input

Description:   Gets the scale the U or I Input

functions for the linearization curve

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

mode

Variable Type       int

Control Name:  Mode

Description:   Specifies the scale mode.

Variable Type: Integer

Valid Range:   0 = U Input (Voltage)

1 = I Input (Current)

u1I1

Variable Type       double (passed by reference)

Control Name:  U1/I1

Description:   Gets the U1 or I1 value

for scale t1

Variable Type: Double

Default:       0.0

Valid Range:   0 - 10V  => U-Input (Voltage)

0 - 20mA => I-Input (Current)

u2I2

Variable Type       double (passed by reference)

Control Name:  U2/I2

Description:   Gets U2 or I2 value

for scale t1

Variable Type: Double

Default:       0.0

Valid Range:   0 - 10V  => U-Input (Voltage)

0 - 20mA => I-Input (Current)

t1

Variable Type       double (passed by reference)

Control Name:  t1

Description:   Gets t1 value

for scale of U1 or I1

Variable Type: Double

Default:       0.0

Valid Range:   -200 - 800 C

t2

Variable Type       double (passed by reference)

Control Name:  t2

Description:   Gets the t2 value

for scale of U2 or I2

Variable Type: Double

Default:       0.0

Valid Range:   -200 - 800 C

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Mode

-3   Invalid U1 or I1

-4   Invalid U2 or I2

-5   Invalid t1

-6   Invalid t2

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_get_config_temp_coefficient

int bp2329_get_config_temp_coefficient (int instrumentID,

int number_ofCoeff,

int *coeffValueppmk);

Purpose

Panel Name:    Get Config Temperature Coefficients

Descrpition:   Gets one of the ten temperature

coefficents.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

number_ofCoeff

Variable Type       int

Control Name:  Number of coefficient

Description:   This parameter selects one of ten

possible temperature coefficients

Variable Type: Integer

Valid Range:   1 to 10

coeffValueppmk

Variable Type       int (passed by reference)

Control Name:  Value of coefficient

Description:   This parameter gets the

value of the desired temperature

coefficients.The unit is ppm/k.

Variable Type: Integer

Valid Range:   -99999 - 99999

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Number of TC

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_config_temp_compensation

int bp2329_get_config_temp_compensation (int instrumentID,

int *compensationState,

int *compensationType,

double *referenceTemperature,

double *manualTemperature);

Purpose

Panel Name:    Get Config Temp. Compensation

Descrpition:

- Gets the state

- Gets the manual temperature

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

compensationState

Variable Type       int (passed by reference)

Control Name:  Compensation State

Description:   This parameter gets the state of the

temperature compensation.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

compensationType

Variable Type       int (passed by reference)

Control Name:  Compensation Type

Description:   This parameter gets the type

temperature compensation.

Variable Type: Integer

Valid Range:   0 = Manualtemperature compensation

1 = PT100 Auto temperature compensation

2 = U-Input Auto temperature compensation

3 = I-Input Auto temperature compensation

referenceTemperature

Variable Type       double (passed by reference)

Control Name:  Manual Temperature

Description:   This parameter gets the manually

temperature value

Variable Type: Double

manualTemperature

Variable Type       double (passed by reference)

Control Name:  Manual Temperature

Description:   This parameter gets the manually

temperature value

Variable Type: Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid State

-3     Invalid Type

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_datalogger_block_size

int bp2329_get_datalogger_block_size (int instrumentID, int blockNo,

int *blockSize);

Purpose

Panel Name:    Get Datalogger Block Size

Description:   Gets the datalogger block size

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to configure

Variable Type: Integer

Valid Range:   0 - 31

blockSize

Variable Type       int (passed by reference)

Control Name:  Block Size

Description:   Gets the size of the block.

The hole memory size of the device

if 20000. If define one block with

20000, you have no memory free to

define another block.

Variable Type: Integer

Valid Range:   0 - 20000

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

-3   Invalid Block size

-4   Block name to long

-5   Invalid Filter

-6   Invalid X-Value

-7   Invalid hours

-8   Invalid minutes

-9   Invalid secondes

-10  Invalid Delta

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_datalogger_filter

int bp2329_get_datalogger_filter (int instrumentID, int blockNo,

int *filter, int *XValue, int *HH,

int *MM, int *SS, double *delta);

Purpose

Panel Name:    Get Config Datalogger

Description:   Gets the datalogger modes and parameter

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to configure

Variable Type: Integer

Valid Range:   0 - 31

filter

Variable Type       int (passed by reference)

Control Name:  Filter

Description:   Gets the filter of the block.

Variable Type: Integer

Valid Range:   0 = All values are logged

1 = Only good values are logged

2 = Only bad values are logged

3 = Only the x-values values are logged

4 = Timer inteval for logging

5 = The value will be logged, if

the difference to the previous

value is greather than then

delta value.

XValue

Variable Type       int (passed by reference)

Control Name:  X-Value

Description:   Gets the X-value .The interval

in witch the value is regularly

stored.

Variable Type: Integer

Valid Range:   2 - 9999

HH

Variable Type       int (passed by reference)

Control Name:  Time HH

Description:   Gets the hours for the time.

Variable Type: Integer

Valid Range:   0 -99

MM

Variable Type       int (passed by reference)

Control Name:  Timer Interval MM

Description:   Gets the minutes for the time.

Variable Type: Integer

Valid Range:   0 -59

SS

Variable Type       int (passed by reference)

Control Name:  Timer Interval SS

Description:   Gets the seconds for the time.

Variable Type: Integer

Valid Range:   0 - 59

delta

Variable Type       double (passed by reference)

Control Name:  Delta Value

Description:   Gets the delta value of the

block.

Note :         The value will be logged, if

the difference to the previous

value is greather than then

delta value.

Variable Type: double

Valid Range:   0.01mOhm - 200 KOhm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

-3   Invalid Block size

-4   Block name to long

-5   Invalid Filter

-6   Invalid X-Value

-7   Invalid hours

-8   Invalid minutes

-9   Invalid secondes

-10  Invalid Delta

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_datalogger_name

int bp2329_get_datalogger_name (int instrumentID, int blockNo,

char name[]);

Purpose

Panel Name:    Get Config Datalogger

Description:   Gets the datalogger name

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to configure

Variable Type: Integer

Valid Range:   0 - 31

name

Variable Type       char []

Control Name:  Block Name

Description:   Specificies the name of the block.

Variable Type: String (max. Length 10 Byte)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

-3   Invalid Block size

-4   Block name to long

-5   Invalid Filter

-6   Invalid X-Value

-7   Invalid hours

-8   Invalid minutes

-9   Invalid secondes

-10  Invalid Delta

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_datalogger_state

int bp2329_get_datalogger_state (int instrumentID,

int *datenloggerState);

Purpose

Panel Name:    Get Datalogger State

Description:   Gets the datalogger state ON/OFF.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

datenloggerState

Variable Type       int (passed by reference)

Control Name:  State

Description:   Gets the datalogger state

Variable Type: Integer

Valid Range:   0 = Off

1 = On

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_envelope_state

int bp2329_get_envelope_state (int instrumentID, int *state);

Purpose

Panel Name:    Get Envelope State

Description:   Gets the Envelope State

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

state

Variable Type       int (passed by reference)

Control Name:  State

Description:   Getsthe State is On the envelope

acquistition

Variable Type: Integer

Valid Range:   0 = Off

1 = On

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_get_temp_value

int bp2329_get_temp_value (int instrumentID, double *temperatureValue);

Purpose

Panel Name:    Get Config Temp. Compensation

Descrpition:

- Gets the state

- Gets the manual temperature

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

temperatureValue

Variable Type       double (passed by reference)

Control Name:  Manual Temperature

Description:   This parameter gets the manually

temperature value

Variable Type: Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid State

-3     Invalid Type

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_init_gpib

int bp2329_init_gpib (int address, int timeout, int reset,

char instrumentIDString[], int *instrumentID);

Purpose

Panel Name:    Initialize Gpib

Description:   Initializes the instrument by:

1) finding an unused device in the device table

using the supplied GPIB address

2) device clear and RESET  (if desired)

3) returns the ID-String of the instrument

4) returns an instrument ID

This is a multi-instrument driver, allowing

multiple Resistomat 2304+2305 instruments to be

controlled.

During the initialization, the returned ID is

based on the GPIB address.  This ID is used in

all of the other functions to distinguish

between multiple Resistomat 2304+2305

instruments.

Parameter List

address

Variable Type       int

Control Name:  Address

Description:   Specifies the address on the GPIB

to be used to communicate with the instrument.

This value must correspond to the current

internal setting of the instrument.

Variable Type: Integer

Valid Range:   0 to 30

timeout

Variable Type       int

Control Name:   Timeout (seconds)

Description:    Specifies  the timeout limit for

I/O operations.

Variable Type:  Integer

Valid Range:    0 - 5

0 = 100msec

1 = 300msec

2 = 1sec

3 = 3sec

4 = 10sec

5 = 30sec

Default:        3sec

reset

Variable Type       int

Control Name:  Reset

Description:   If reset is on, a device clear and

power on reset is done during initialization.

Variable Type: Integer

Valid Range:   0 to 1

0  off

1  on

instrumentIDString

Variable Type       char []

Control Name:  Instrument ID String

Description:   Returns the ID-String of the

instrument.

Variable Type: String

instrumentID

Variable Type       int (passed by reference)

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid GPIB Address

-2     Invalid Reset

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

________________________________________________________________________________

bp2329_init_serial

int bp2329_init_serial (int COMPort, int baudRate,

double timeout_seconds,

char instrumentIDString[], int reset,

int *instrumentID);

Purpose

Panel Name:    Initialize Serial

Description:   Initializes the instrument by:

1) finding an unused device in the device table

using the supplied serial group and user

address

2) device clear and RESET  (if desired)

3) returns the ID-String of the instrument

4) returns an instrument ID

This is a multi-instrument driver, allowing

multiple Resistomat 2304-2305 instruments to be

controlled.

During the initialization, the returned ID is

based on the serial group and user address.

This ID is used in all of the other functions

to distinguish between multiple Resistomat 2304+

2305 instruments.

Parameter List

COMPort

Variable Type       int

Control Name:   COM-Port

Description:    Specifies the COM-Port to be

used to communicate with the

instrument.

Variable Type:  Integer

Valid Values:   1 - COM1   (=> Adress 0x3f8 IRQ4)

2 - COM2   (=> Adress 0x2f8 IRQ3)

3 - COM3   (=> Adress 0x3e8 IRQ4)

4 - COM4   (=> Adress 0x2e8 IRQ3)

Default:        1

baudRate

Variable Type       int

Control Name:   Baud Rate

Description:    Specifies the baud rate of the

RS232-Port to communicate with

the instrument. This value must

correspond to the current

internal setting of the

instrument.

Variable Type:  Integer

Valid Values:   1200,2400,4800,9600,19200 or 38400

baud

Default:        9600

timeout_seconds

Variable Type       double

Control Name:   Time-out (seconds)

Description:    Specifies  the time-out limit for

I/O operations.

A  value  of  0.0  will  disable

time-out altogether.

Variable Type:  Double

Valid Range:    0.060 to 100.0 seconds

Default:        1.0

instrumentIDString

Variable Type       char []

Control Name:  Instrument ID String

Description:   Returns the ID-String of the

instrument.

Variable Type: String

reset

Variable Type       int

Control Name:  Reset

Description:   If reset is on, a device clear and

power on reset is done during initialization.

Variable Type: Integer

Valid Range:   0 to 1

0 = off

1 = on

instrumentID

Variable Type       int (passed by reference)

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   COM-Port is out of range

-2   Baud Rate is out of range

-3   Time-out is out of range

-4   Block Check parameter is out of range

-5   Group Address is out of range

-6   User Address is out of range

-7   Reset Device is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_local_gpib

int bp2329_local_gpib (int instrumentID);

Purpose

Panel Name:    Local (GPIB only)

Description:   Sets the  temporarily in

local mode.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

________________________________________________________________________________

bp2329_meas_read

int bp2329_meas_read (int instrumentID, double *measuredValue,

char unit[]);

Purpose

Panel Name:    Measure Read

Descrpition:   Revieves measured values while

a measurement is done

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

measuredValue

Variable Type       double (passed by reference)

The control return the measurement value

Variable Type : double

unit

Variable Type       char []

The control return the measurement unit

Variable Type : String (Length=8Byte)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT => No Value available

302   There ist's no measurement started

303   Timeout for wait of RQS Byte

304   GPIB Error for read the RQS Byte

System Device Errors :

----------------------------------------------

310   Temperature Compensation Error

311   U-Cable Error

312   I-Oszillation Error

313   Under Load Error

314   Over Load Error

315   I-Over Error

316   Meas.Amplifier U-Overrange Error

317   Meas.Amplifier I-Overrange Error

________________________________________________________________________________

bp2329_meas_read_string

int bp2329_meas_read_string (int instrumentID, char measuredValue[]);

Purpose

Panel Name:    Measure Read

Descrpition:   Revieves measured values while

a measurement is done

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

measuredValue

Variable Type       char []

The control return the measurement value

Variable Type : double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT => No Value available

302   There ist's no measurement started

303   Timeout for wait of RQS Byte

304   GPIB Error for read the RQS Byte

System Device Errors :

----------------------------------------------

310   Temperature Compensation Error

311   U-Cable Error

312   I-Oszillation Error

313   Under Load Error

314   Over Load Error

315   I-Over Error

316   Meas.Amplifier U-Overrange Error

317   Meas.Amplifier I-Overrange Error

________________________________________________________________________________

bp2329_meas_start

int bp2329_meas_start (int instrumentID);

Purpose

Panel Name:    Measure Start

Descrpition:   Start the measurement

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_meas_stop

int bp2329_meas_stop (int instrumentID);

Purpose

Panel Name:    Measure Stop

Descrpition:   Stop the measurement

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_printer_counter_reset

int bp2329_printer_counter_reset (int instrumentID);

Purpose

Panel Name:    Reset Printer Counter

Description:   Reset the printer counter

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_receive_scpi

int bp2329_receive_scpi (int instrumentID, char sendSCPIMessage[],

char answer_fromSlave[]);

Purpose

Panel Name:   Receive a SCPI-Message

Descrpition:  Send "Query form" from SCPI-

Message and receive a answers

=> Polling

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

sendSCPIMessage

Variable Type       char []

Control Name:  Send SCPI-Message

Description:   This control is a string parameter

You must input Query-Form from

the SCPI-message without the

LineFeed termination.

Example      : ":SENS:CORR:STATE?"

Variable Type: String

answer_fromSlave

Variable Type       char []

Control Name:  Answer from Slave

Description:   This control return a string from

the slave.

Variable Type: String (Lenght => see the 2304

Manual SCPI-Commands)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_send_scpi

int bp2329_send_scpi (int instrumentID, char sendSCPIMessage[]);

Purpose

Panel Name:   Send a SCPI-Message

Descrpition:  Send every SCPI-Message to a

slave

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

sendSCPIMessage

Variable Type       char []

Control Name:  Send SCPI-Message

Description:   This control is a string parameter

You must input the SCPI-message

without the LineFeed termination.

Example      : ":SENS:CORR:STATE ON"

Variable Type: String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_active_datalogger_block

int bp2329_set_active_datalogger_block (int instrumentID,

int selectMode, int blockNo,

char name[]);

Purpose

Panel Name:    Set Active Datalogger Block

Description:   Setect Active Datalogger Block

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

selectMode

Variable Type       int

Control Name:  Select Mode

Description:   Switch the selecte mode, do you

want to set the active block

Variable Type: Integer

Valid Range:   0 = Set active by block no.

1 = Set active by block name

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to set as active block.

Variable Type: Integer

Valid Range:   0 - 31

name

Variable Type       char []

Control Name:  Block Name

Description:   Specificies the name of the block.

Variable Type: String (max. Length 10 Byte)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_active_temp_coefficient

int bp2329_set_active_temp_coefficient (int instrumentID,

int number_ofCoeff);

Purpose

Panel Name:    Set Avtive Temperature Coefficients

Descrpition:   Sets one of the ten temperature

coefficents.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

number_ofCoeff

Variable Type       int

Control Name:  Number of coefficient

Description:   This parameter selects one of ten

possible temperature coefficients

Variable Type: Integer

Valid Range:   1 to 10

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Number of TC

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_2_limit_comparator

int bp2329_set_config_2_limit_comparator (int instrumentID,

double lowerLimit,

double upperLimit);

Purpose

Panel Name:    Set Two Limit Comparator

Description:   Configure the two limit comparator

functions

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

lowerLimit

Variable Type       double

Control Name:  Lower Limit

Description:   Specifies the lower limit of the

two limit comparator

Variable Type: Double

Default:       0.0

upperLimit

Variable Type       double

Control Name:  Upper Limit

Description:   Specifies the upper limit of the

two limit comparator

Variable Type: Double

Default:       0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_set_config_4_limit_comparator

int bp2329_set_config_4_limit_comparator (int instrumentID,

double limit1, double limit2,

double limit3, double limit4);

Purpose

Panel Name:    Set Four Limit Comparator

Description:   Configure the four limit

comparator functions

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

limit1

Variable Type       double

Control Name:  Limit 1

Description:   Specifies the limit 1 of the

four limit comparator

Variable Type: Double

Default:       0.0

limit2

Variable Type       double

Control Name:  Limit2

Description:   Specifies the limit 2 of the

four limit comparator

Variable Type: Double

Default:       0.0

limit3

Variable Type       double

Control Name:  Limit 3

Description:   Specifies the limit 3 of the

four limit comparator

Variable Type: Double

Default:       0.0

limit4

Variable Type       double

Control Name:  Limit4

Description:   Specifies the limit 4 of the

four limit comparator

Variable Type: Double

Default:       0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_set_config_auto_range

int bp2329_set_config_auto_range (int instrumentID, int lowerAutoRange,

int upperAutoRange);

Purpose

Panel Name:    Set Config Auto Range

Description:   Configure the instrument auto range

function in

following ways:

- Sets the upper and lower range for

the auto range function of the device

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

lowerAutoRange

Variable Type       int

Control Name:  Lower Auto Range

Description:   This parameter selects the lowest

range in automatic measurement

range selection

Variable Type: Integer

Valid Range:   1 = 200 mOhm

2 = 2    Ohm

3 = 20   Ohm

4 = 200  Ohm

5 = 2   KOhm

6 = 20  KOhm

7 = 200 KOhm

upperAutoRange

Variable Type       int

Control Name:  Upper Auto Range

Description:   This parameter selects the highest

range in automatic measurement

range selection

Variable Type: Integer

Valid Range:   1 = 200 mOhm

2 = 2    Ohm

3 = 20   Ohm

4 = 200  Ohm

5 = 2   KOhm

6 = 20  KOhm

7 = 200 KOhm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid lower Range

-3   Invalid upper range

-4   upper range <= lower range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_comparator_mode

int bp2329_set_config_comparator_mode (int instrumentID, int state,

int relaisMode, int onError);

Purpose

Panel Name:    Set Comparator Mode

Description:   Sets the Comparator State ON/OFF

Config the relais mode

Config the comparator fault mode

(On Error)

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

state

Variable Type       int

Control Name:  State

Description:   If the State is On the comparator

is active

Variable Type: Integer

Valid Range:   0 = Off

1 = On

relaisMode

Variable Type       int

Control Name:  Relais Mode

Description:   If the relais mode is On, the

comparator relais are active

Variable Type: Integer

Valid Range:   0 = Off

1 = On

onError

Variable Type       int

Control Name:  On Error

Description:   If the On Error is upper,the

comparator show the ">" - token

by an error

Variable Type: Integer

Valid Range:   0 = None

1 = Upper ">"-token

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

-3   Invalid relais mode

-4   Invalid on error

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_device_access

int bp2329_set_config_device_access (int instrumentID, int measureMenu,

int displayMenu, int printerMenu,

int tempCompMenu, int contrastMenu,

int storeConfigMenu,

int interfaceMenu,

int displayStateMenu,

int auxInOutMenu, int scaleMenu,

int dataloggerMenu,

int calibrationMenu, int dateMenu,

int fastStartMenu,

int comparaConfMenu,

int comparaEvalMenu,

int envelopeConfMenu,

int envelopeEvalMenu,

int manRangeMenu);

Purpose

Panel Name:    Set Config Device Access

Description:   Configure the access to the device

menues.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

measureMenu

Variable Type       int

Control Name:  Measure Menu

Description:   This parameter activate or

deactivate the access to the

measure menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

displayMenu

Variable Type       int

Control Name:  Display Menu

Description:   This parameter activate or

deactivate the access to the

display menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

printerMenu

Variable Type       int

Control Name:  Printer Menu

Description:   This parameter activate or

deactivate the access to the

Printer menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

tempCompMenu

Variable Type       int

Control Name:  Temp. Comp. Menu

Description:   This parameter activate or

deactivate the access to the

temperature compensation menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

contrastMenu

Variable Type       int

Control Name:  Contrast Menu

Description:   This parameter activate or

deactivate the access to the

contrast menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

storeConfigMenu

Variable Type       int

Control Name:  Store Config. Menu

Description:   This parameter activate or

deactivate the access to the

store config. menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

interfaceMenu

Variable Type       int

Control Name:  Interface Menu

Description:   This parameter activate or

deactivate the access to the

interface menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

displayStateMenu

Variable Type       int

Control Name:  Display State Menu

Description:   This parameter activate or

deactivate the access to the

display state menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

auxInOutMenu

Variable Type       int

Control Name:  Aux. In/Out Menu

Description:   This parameter activate or

deactivate the access to the

Aux. Input/Output menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

scaleMenu

Variable Type       int

Control Name:  Scale Menu

Description:   This parameter activate or

deactivate the access to the

scale menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

dataloggerMenu

Variable Type       int

Control Name:  Datalogger Menu

Description:   This parameter activate or

deactivate the access to the

datalogger menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

calibrationMenu

Variable Type       int

Control Name:  Calibration Menu

Description:   This parameter activate or

deactivate the access to the

Calibration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

dateMenu

Variable Type       int

Control Name:  Date Menu

Description:   This parameter activate or

deactivate the access to the

date menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

fastStartMenu

Variable Type       int

Control Name:  Fast Start Menu

Description:   This parameter activate or

deactivate the access to the

fast start menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

comparaConfMenu

Variable Type       int

Control Name:  Comparator Config Menu

Description:   This parameter activate or

deactivate the access to the

comparator config menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

comparaEvalMenu

Variable Type       int

Control Name:  Comparator Eval. Menu

Description:   This parameter activate or

deactivate the access to the

comparator eval. menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

envelopeConfMenu

Variable Type       int

Control Name:  Envelope Config. Menu

Description:   This parameter activate or

deactivate the access to the

envelope config (Envelope Start) menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

envelopeEvalMenu

Variable Type       int

Control Name:  Envelope Eval. Menu

Description:   This parameter activates or

deactivates the access to the

envelope evalation menu(Reset Button).

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

manRangeMenu

Variable Type       int

Control Name:  Manual Range Menu

Description:   This parameter activate or

deactivate the access to the

manual range menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid meaure menu

-3     Invalid display menu

-4     Invalid printer menu

-5     Invalid temp.comp menu

-6     Invalid contrast menu

-7     Invalid store config. menu

-8     Invalid interface menu

-9     Invalid display state menu

-10     Invalid aux. input/output menu

-11     Invalid scale menu

-12     Invalid datalogger menu

-13     Invalid calibration menu

-14     Invalid date menu

-15     Invalid fast start menu

-16     Invalid comparator config. menu

-17     Invalid comparator eval. menu

-18     Invalid envelope config. menu

-19     Invalid envelope eval. menu

-20     Invalid manuall range menu

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_display_contrast

int bp2329_set_config_display_contrast (int instrumentID,

double displayContrast);

Purpose

Panel Name:    Set Config Display Contrast

Descrpition:   Set the display contrast

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

displayContrast

Variable Type       double

Control Name:  Display Contrast

Description:   This parameter is meant for

adjusting the display contrast.

Variable Type: Double

Valid Range:   0.0 - 1.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Contrast Value

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_display_mode

int bp2329_set_config_display_mode (int instrumentID, int displayMode,

double deltaNominalValue,

int resolution);

Purpose

Panel Name:    Set Config Display Mode

Description:   Configure the instrument display mode

following ways:

- Selects the display mode

- Selects the display resolution

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

displayMode

Variable Type       int

Control Name:  Display Mode

Description:   This parameter select the display mode.

Variable Type: Integer

Valid Range:   0 = Ohm

1 = Delta %

2 = Comparator

deltaNominalValue

Variable Type       double

Control Name:  Display Nominal Value

Description:   This parameter inputs the nominal value

from the delta% display mode.

Variable Type: double

resolution

Variable Type       int

Control Name:  Resolution

Description:   This parameter select the resolution

of the device.

Variable Type: Integer

Valid Range:   0 = 2000 Digits

1 = 20000 Digits

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid display mode

-3     Invalid resolution

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_display_options

int bp2329_set_config_display_options (int instrumentID,

int rangeOption,

int currentOption,

int temperatureOption,

int sampleOption,

int triggerOption,

int autoManOption,

int loadOption,

int measureOption);

Purpose

Panel Name:    Set Config Display Options

Descrpition:   Config the display options of the

device

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

rangeOption

Variable Type       int

Control Name:  Range Option

Description:   This parameter activate or

deactivate the range function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

currentOption

Variable Type       int

Control Name:  Current Option

Description:   This parameter activate or

deactivate the current function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

temperatureOption

Variable Type       int

Control Name:  Temperature Option

Description:   This parameter activate or

deactivate the temperature function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

sampleOption

Variable Type       int

Control Name:  Sample Option

Description:   This parameter activate or

deactivate the sample function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

triggerOption

Variable Type       int

Control Name:  Trigger Option

Description:   This parameter activate or

deactivate the trigger function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

autoManOption

Variable Type       int

Control Name:  Auto/Man Option

Description:   This parameter activate or

deactivate the Auto/Man function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

loadOption

Variable Type       int

Control Name:  Load Option

Description:   This parameter activate or

deactivate the load function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

measureOption

Variable Type       int

Control Name:  Measure Option

Description:   This parameter activate or

deactivate the measure function

on the device display.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid range option

-3     Invalid current option

-4     Invalid temperature option

-5     Invalid sample option

-6     Invalid trigger option

-7     Invalid auto/man option

-8     Invalid load option

-9     Invalid measure option

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_meas

int bp2329_set_config_meas (int instrumentID, int averageMode,

int averageCounts, int loading,

int measurementMode, int _20mVLimit,

int measurementSequence,

int num_ofConvertsCycles);

Purpose

Panel Name:    Set Config Measurement

Description:   Configure the instrument in

following ways:

- Sets the average counts

- Selects the average control

- Selects the loading mode

- Selects the measurement mode

- Selects the 20 mV limit

- Selects the measurement sequence

- Selects the numbers of convert Cycles

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

averageMode

Variable Type       int

Control Name:  Average Control

Description:   This parameter selects the type of

Averaging

Variable Type: Integer

Valid Range:   0 = Moving

1 = Repeat

averageCounts

Variable Type       int

Control Name:  Average Counts

Description:   This parameter inputs the numbers

of average counts

Variable Type: Integer

Valid Range:   1 - 9999

loading

Variable Type       int

Control Name: Loading

Description: This parameter selects the load to

measured (mainly ohmic or mainly

indictive resistance).

Variable Type: Integer

Valid Range:   0 = mainly ohmic

1 = mainly inductive

measurementMode

Variable Type       int

Control Name:  Measurement Mode

Description:   This parameter is meant for

changing between single and

repetition measurements.

Variable Type: Integer

Valid Range:   0 = Single

1 = Repetition

_20mVLimit

Variable Type       int

Control Name:  20 mV Limit

Description:   Selects the range of the 20 mV

Limit.

Variable Type: Integer

Valid Range:   0 = Off

1 = ON

measurementSequence

Variable Type       int

Control Name:  Measurement Sequence

Description:   This parameter selects then measurement

sequence.

Variable Type: Integer

Valid Range:   0 = Ref-Comp

(Zero measuremnet with current = 0)

1 = None Comp

(Measurement without zero measurement)

2 = Single Comp

(Measurement with one zero measurement)

3 = Standard

(Measurement with toggle between zero

measurement and realy measurement)

num_ofConvertsCycles

Variable Type       int

Control Name:  Number of Converts

Description:   This parameter select the number of converts.

Variable Type: Integer

Valid Range:   0 = Minimum

1 = Medium

2 = Maximum

3 = Standard

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Average Control

-3   Invalid Average Counts

-4   Invalid Loading Mode

-5   Invalid Measurement Mode

-6   Invalid 20 mV State

-7   Invalid Measurement Sequence

-8   Invalid Numbers of Convert Cycles

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_memory

int bp2329_set_config_memory (int instrumentID, int memoryNo);

Purpose

Panel Name:    Store Memory No.

Description:   Store the actual Parameters to

the selected memory of the

instrument

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

memoryNo

Variable Type       int

Control Name:   Memory No.

Description:    Specifies the Memory No.

Variable Type:  integer

Valid Range:    0 to 32

Default:        0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Memory No.

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

320   Datalogger is empty

________________________________________________________________________________

bp2329_set_config_memory_label

int bp2329_set_config_memory_label (int instrumentID, int memoryNo,

char memoryLabel[]);

Purpose

Panel Name:    Set Memory Labels

Description:   Store the memory label to the

specificies memory No.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

memoryNo

Variable Type       int

Control Name:  Memory No.State

Description:   Specificies the memory No. to

store the the label text

Variable Type: Integer

Valid Range:   0 -  32

memoryLabel

Variable Type       char []

Control Name:  Memory Label

Description:   Specifices the memory label.

Variable Type: String

Valid Range:   max. 10 Byte

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

-3   Memory label to long

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_printer_mode

int bp2329_set_config_printer_mode (int instrumentID,

int printTemperature,

int printCounter, int printLabel,

int printTime, int printDate,

int measValuesHeader,

char labelText[], int HH, int MM,

int SS);

Purpose

Panel Name:    Set Config Printer Mode

Description:   Configure the printer options

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

printTemperature

Variable Type       int

Control Name:  Print Temperature

Description:   Switch the print temperature function

ON or OFF.

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printCounter

Variable Type       int

Control Name:  Print Counter

Description:   Switch the print counter function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printLabel

Variable Type       int

Control Name:  Print Label

Description:   Switch the print label function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printTime

Variable Type       int

Control Name:  Print Time

Description:   Switch the print time function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

printDate

Variable Type       int

Control Name:  Print Date

Description:   Switch the print date function

ON or OFF

Variable Type: Integer

Valid Range:   0 = Off

1 = On

measValuesHeader

Variable Type       int

Control Name:  Meas.Values/Header

Description:   Specifices how many measure values

do want print per header.

Variable Type: Integer

Valid Range:   1 - 1000

labelText

Variable Type       char []

Control Name:  Label text

Description:   Specifices the label text for the

printer.

Variable Type: String

Valid Range:   max. 10 Byte

HH

Variable Type       int

Control Name:  Timer Interval HH

Description:   Specifices the hours for the timer

interval.

Variable Type: Integer

Valid Range:   0 -99

MM

Variable Type       int

Control Name:  Timer Interval MM

Description:   Specifices the minutes for the timer

interval.

Variable Type: Integer

Valid Range:   0 -59

SS

Variable Type       int

Control Name:  Timer Interval SS

Description:   Specifices the seconds for the timer

interval.

Variable Type: Integer

Valid Range:   0 - 59

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Print Temperature

-3   Invalid Print Counter

-4   Invalid Print Label

-5   Invalid Print Time

-6   Invalid Print Date

-7   Invalid Meas.Values/Header

-8   Invalid Label Text (to long)

-9   Invalid Timer Interval HH

-10   Invalid Timer Interval MM

-11   Invalid Timer Interval SS

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_printer_state

int bp2329_set_config_printer_state (int instrumentID, int printerState);

Purpose

Panel Name:    Set Config Printer State

Description:   Start or stop the printer

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

printerState

Variable Type       int

Control Name:  Printer State

Description:   If the printer started or stopped

Variable Type: Integer

Valid Range:   0 = Stop

1 = Start

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

-3   Invalid relais mode

-4   Invalid on error

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_range

int bp2329_set_config_range (int instrumentID, int range);

Purpose

Panel Name:    Set Config Range

Description:   Configure the instrument range in

following ways:

- Selects the Auto range function

or sets the range manuall

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

range

Variable Type       int

Control Name:  Range

Description:   This parameter activates or de-

activates automatic measuremnet

range selection and selects the

desired measurement range.

Variable Type: Integer

Valid Range:   0 = Auto

1 = 200 mOhm

2 = 2    Ohm

3 = 20   Ohm

4 = 200  Ohm

5 = 2   KOhm

6 = 20  KOhm

7 = 200 KOhm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_scale_pt100

int bp2329_set_config_scale_pt100 (int instrumentID, double r0,

double coefficienta,

double coefficientb);

Purpose

Panel Name:    Set Config Scale PT100

Description:   Scale the PT100 function

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

r0

Variable Type       double

Control Name:  R0

Description:   Specifies the R0

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Variable Type: Double

Default:       0.0

coefficienta

Variable Type       double

Control Name:  Coefficient "a"

Description:   Specifies the coefficient "a".

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Variable Type: Double

Default:       0.0

coefficientb

Variable Type       double

Control Name:  Coefficient "b"

Description:   Specifies the coefficient "b".

Formula :      Rt = R0 *(1 + "a" * t + "b" * t^2 )

Variable Type: Double

Default:       0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_set_config_scale_ui_input

int bp2329_set_config_scale_ui_input (int instrumentID, int mode,

double u1I1, double u2I2,

double t1, double t2);

Purpose

Panel Name:    Set Config Scale U/I Input

Description:   Scale the U or I Input functions

for the linearization curve

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

mode

Variable Type       int

Control Name:  Mode

Description:   Specifies the scale mode.

Variable Type: Integer

Valid Range:   0 = U Input (Voltage)

1 = I Input (Current)

u1I1

Variable Type       double

Control Name:  U1/I1

Description:   Specifies the U1 or I1 value

for scale t1

Variable Type: Double

Default:       0.0

Valid Range:   0 - 10V  => U-Input (Voltage)

0 - 20mA => I-Input (Current)

u2I2

Variable Type       double

Control Name:  U2/I2

Description:   Specifies the U2 or I2 value

for scale t1

Variable Type: Double

Default:       0.0

Valid Range:   0 - 10V  => U-Input (Voltage)

0 - 20mA => I-Input (Current)

t1

Variable Type       double

Control Name:  t1

Description:   Specifies the t1 value

for scale of U1 or I1

Variable Type: Double

Default:       0.0

Valid Range:   -200 - 800 C

t2

Variable Type       double

Control Name:  t2

Description:   Specifies the t2 value

for scale of U2 or I2

Variable Type: Double

Default:       0.0

Valid Range:   -200 - 800 C

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Mode

-3   Invalid U1 or I1

-4   Invalid U2 or I2

-5   Invalid t1

-6   Invalid t2

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

330   Invalid Comparator Datas

________________________________________________________________________________

bp2329_set_config_temp_coefficient

int bp2329_set_config_temp_coefficient (int instrumentID,

int number_ofCoeff,

int coeffValueppmk);

Purpose

Panel Name:    Set Config Temperature Coefficients

Descrpition:   Config one of the ten temperature

coefficents.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

number_ofCoeff

Variable Type       int

Control Name:  Number of coefficient

Description:   This parameter selects one of ten

possible temperature coefficients

Variable Type: Integer

Valid Range:   1 to 10

coeffValueppmk

Variable Type       int

Control Name:  Value of coefficient

Description:   This parameter specifices the

value of the desired temperature

coefficients.The unit is ppm/k.

Variable Type: Integer

Valid Range:   -99999 - 99999

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Number of TC

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_config_temp_compensation

int bp2329_set_config_temp_compensation (int instrumentID,

int compensationState,

int compensationType,

double refernceTemperature,

double manualTemperature);

Purpose

Panel Name:    Set Config Temp. Compensation

Descrpition:   Set the follow parameters:

- toggle the state ON/OFF

- Input the manual temperature

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

compensationState

Variable Type       int

Control Name:  Correction State

Description:   This parameter specifies whether a

temperature compensation si to be

carried out during the resistance

measurement.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

compensationType

Variable Type       int

Control Name:  Correction Type

Description:   This parameter specifies whether a

temperature value in the case of

temperature compensation should be

registered automatically or

entered manually.

Variable Type: Integer

Valid Range:   0 = Manualtemperature compensation

1 = PT100 Auto temperature compensation

2 = U-Input Auto temperature compensation

3 = I-Input Auto temperature compensation

refernceTemperature

Variable Type       double

Control Name:  Manual Temperature

Description:   This parameter specifies the

temperature value in the case of

temperature compensation should be

registered manually.

Variable Type: Double

manualTemperature

Variable Type       double

Control Name:  Manual Temperature

Description:   This parameter specifies the

temperature value in the case of

temperature compensation should be

registered manually.

Variable Type: Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid State

-3     Invalid Type

-4     Invalid Reference Temperature

-5     Invalid Manual Temperature

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_datalogger_block_size

int bp2329_set_datalogger_block_size (int instrumentID, int blockNo,

int blockSize);

Purpose

Panel Name:    Set Datalogger Size

Description:   Sets the datalogger block size

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to configure

Variable Type: Integer

Valid Range:   0 - 31

blockSize

Variable Type       int

Control Name:  Block Size

Description:   Specificies the size of the block.

The hole memory size of the device

if 20000. If define one block with

20000, you have no memory free to

define another block.

Variable Type: Integer

Valid Range:   0 - 20000

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

-3   Invalid Block size

-4   Block name to long

-5   Invalid Filter

-6   Invalid X-Value

-7   Invalid hours

-8   Invalid minutes

-9   Invalid secondes

-10  Invalid Delta

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_datalogger_filter

int bp2329_set_datalogger_filter (int instrumentID, int blockNo,

int filter, int XValue, int HH,

int MM, int SS, double delta);

Purpose

Panel Name:    Set Config Datalogger

Description:   Configure the datalogger modes

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to configure

Variable Type: Integer

Valid Range:   0 - 31

filter

Variable Type       int

Control Name:  Filter

Description:   Specificies the filter of the block.

Variable Type: Integer

Valid Range:   0 = All values are logged

1 = Only good values are logged

2 = Only bad values are logged

3 = Only the x-values values are logged

4 = Timer inteval for logging

5 = The value will be logged, if

the difference to the previous

value is greather than then

delta value.

XValue

Variable Type       int

Control Name:  X-Value

Description:   The X-value defines the interval

in witch the value is regularly

stored.

Variable Type: Integer

Valid Range:   2 - 9999

HH

Variable Type       int

Control Name:  Time HH

Description:   Specifices the hours for the time.

Variable Type: Integer

Valid Range:   0 -99

MM

Variable Type       int

Control Name:  Timer Interval MM

Description:   Specifices the minutes for the time.

Variable Type: Integer

Valid Range:   0 -59

SS

Variable Type       int

Control Name:  Timer Interval SS

Description:   Specifices the seconds for the time.

Variable Type: Integer

Valid Range:   0 - 59

delta

Variable Type       double

Control Name:  Delta Value

Description:   Specificies the delta value of the

block.

Note :         The value will be logged, if

the difference to the previous

value is greather than then

delta value.

Variable Type: double

Valid Range:   0.01mOhm - 200 KOhm

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

-3   Invalid Block size

-4   Block name to long

-5   Invalid Filter

-6   Invalid X-Value

-7   Invalid hours

-8   Invalid minutes

-9   Invalid secondes

-10  Invalid Delta

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_datalogger_name

int bp2329_set_datalogger_name (int instrumentID, int blockNo,

char name[]);

Purpose

Panel Name:    Set Datalogger Name

Description:   Set datalogger Name

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

blockNo

Variable Type       int

Control Name:  Block No.

Description:   Specificies the block no., do you

want to configure

Variable Type: Integer

Valid Range:   0 - 31

name

Variable Type       char []

Control Name:  Block Name

Description:   Specificies the name of the block.

Variable Type: String (max. Length 10 Byte)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Block No.

-3   Invalid Block size

-4   Block name to long

-5   Invalid Filter

-6   Invalid X-Value

-7   Invalid hours

-8   Invalid minutes

-9   Invalid secondes

-10  Invalid Delta

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_datalogger_state

int bp2329_set_datalogger_state (int instrumentID, int datenloggerState);

Purpose

Panel Name:    Set Datalogger State

Description:   Sets the datalogger state ON/OFF.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

datenloggerState

Variable Type       int

Control Name:  State

Description:   Switch the datalogger state is On or Off.

Variable Type: Integer

Valid Range:   0 = Off

1 = On

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_set_envelope_state

int bp2329_set_envelope_state (int instrumentID, int state);

Purpose

Panel Name:    Set Envelope State

Description:   Sets the Envelope State ON/OFF

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

state

Variable Type       int

Control Name:  State

Description:   If the State is On the envelope

acquistition is active

Variable Type: Integer

Valid Range:   0 = Off

1 = On

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_system_date_time

int bp2329_system_date_time (int instrumentID, int HH, int MM, int SS,

int day, int month, int year);

Purpose

Panel Name:   System Date/Time

Description:  Sets the System clock

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

HH

Variable Type       int

Control Name:  System Timer HH

Description:   Specifices the hours for the time

Variable Type: Integer

Valid Range:   0 - 23

MM

Variable Type       int

Control Name:  System Time MM

Description:   Specifices the minutes for the timer

Variable Type: Integer

Valid Range:   0 -59

SS

Variable Type       int

Control Name:  System Time SS

Description:   Specifices the seconds for the timer

Variable Type: Integer

Valid Range:   0 - 59

day

Variable Type       int

Control Name:  System Date (Day)

Description:   Specifices the day for the system

date.

Variable Type: Integer

Valid Range:   0 - 31

month

Variable Type       int

Control Name:  System Date (Month)

Description:   Specifices the month for the system

date.

Variable Type: Integer

Valid Range:   1 - 12

year

Variable Type       int

Control Name:  System Date (Year) YYYY

Description:   Specifices the year (YYYY) for the

system date .

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_system_error

int bp2329_system_error (int instrumentID);

Purpose

Panel Name:   System Error

Description:  Reads the system error of the

device

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_system_get_register

int bp2329_system_get_register (int instrumentID, int *bit3, int *bit2,

int *bit1, int *bit0);

Purpose

Panel Name:   Get Register

Descrpition:  Get the input register status

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

bit3

Variable Type       int (passed by reference)

Control Name:  Bit 3

Description:   Return the value of Bit 3 from

I/O register.

Variable Type: Integer

bit2

Variable Type       int (passed by reference)

Control Name:  Bit 2

Description:   Return the value of Bit 2 from

I/O register.

Variable Type: Integer

bit1

Variable Type       int (passed by reference)

Control Name:  Bit 1

Description:   Return the value of 1 from

I/O register.

Variable Type: Integer

bit0

Variable Type       int (passed by reference)

Control Name:  Bit 0

Description:   Return the value of Bit 0

from I/O register.

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Pin 3  Value

-3     Invalid Pin 21  Value

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_system_keyboard

int bp2329_system_keyboard (int instrumentID, int keyboard);

Purpose

Panel Name:   System Keybroad

Description:  Controls the keybroad of the device

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

keyboard

Variable Type       int

Control Name:  Keyboard

Description:   This parameter locks the keyboard

on the front panel of the

instrument

Variable Type: Integer

Valid Range:   0  = Unlocked

1  = Locked

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Value

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp2329_system_set_register

int bp2329_system_set_register (int instrumentID, int bit2, int bit1,

int bit0);

Purpose

Panel Name:   Set Output Register

Descrpition:  Set/Reset the output register

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

bit2

Variable Type       int

Control Name:  Bit 2

Description:   With this parameter can be set or

reset the Bit 2 of I/O register.

Variable Type: Integer

Valid Range:   0  = OFF

1  = ON

bit1

Variable Type       int

Control Name:  Bit 1

Description:   With this parameter can be set or

reset the Bit 1 of I/O register.

Variable Type: Integer

Valid Range:   0  = OFF

1  = ON

bit0

Variable Type       int

Control Name:  Bit 0

Description:   With this parameter can be set or

reset the Bit 0 of I/O register.

Variable Type: Integer

Valid Range:   0  = OFF

1  = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Bit 2 value

-3     Invalid Bit 1 value

-4     Invalid Bit 0 value

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT