Firmware version V1.55 for the 8650A power meter adds the following SCPI commands for modulated measurements.

| SCPI COMMAND | EXISTING NATIVE COMMAND | DESCRIPTION |
| --- | --- | --- |
| SENSe<sensor 1 or 2>:CONFig:BAP:BEEXclude space <samples> (Where <samples> is an integer in the range 0 to 123) Example: SENS1:CONF:BAP:BEEX 2 Set the burst end exclude to 2, i.e. 54μs. | BSPE | Set number of burst end exclude samples. The timing for each sample is approximately 27μs. |
| SENSe<sensor 1 or 2>:CONFig:BAP:BEEXclude? | MEAS [A? or B?] | Read the number of burst end exclude samples. |
| SENSe<sensor 1 or 2>:CONFig:BAP:BSEXclude space <samples> (Where <samples> is an integer in the range 0 to 1565.) Example: SENS1:CONF:BAP:BSEX 1 Set the burst start exclude to 1, i.e. 27μs | BSTE | Set number of burst start exclude samples. The timing for each sample is approximately 27μs.. |
| SENSe<sensor 1 or 2>:CONFig:BAP:BSEXclude? | MEAS [A? or B?] | Read number of burst start exclude samples. |
| SENSe<sensor 1 or 2>:CONFig:BAP:BDTolerance space <dropout> (Where <dropout> is the dropout time in milliseconds in the range 0 to 74 and with a resolution of 0.1 ms. The system selects the nearest value from a series of discrete values.) Example: SENS1:CONF:BAP:BDT 0.054 Set the burst dropout tolerance to 54μs. | BTDP | Set burst dropout tolerance duration in milliseconds |
| SENSe<sensor 1 or 2>:CONFig:BAP:BDTolerance? | MEAS [A? or B?] | Read the current burst dropout tolerance duration. |
| CALCulate<channel 1 to 4>CFACtor:STATe space <On ¦ Off> Examples: CALC1:CFAC:STAT ON Enable crest factor measurements. CALC1:CFAC:STAT? Read the current setting. | CR1 / CR0 | Enable / disable crest factor measurements. Sending ‘On’ when crest factor is already enabled causes the value to be reset. |
| CALCulate<channel 1 to 4>CFACtor:DATa? | CRF | Read the crest factor measurement. |
| CALCulate<channel 1 to 4>PHOLd:STATe space <On ¦ Off> Examples: CALC1:PHOL:STAT ON Enable peak hold measurements. CALC1:CFAC:STAT? Read the current setting. | PH1 / PH0 | Enable / disable peak hold measurements. Sending ‘On’ when peak hold is already enabled causes the value to be reset. |
| CALCulate<channel 1 to 4>PHOLd:DATa? | PKH | Read the peak hold measurement. |
| SENSe<sensor 1 or 2>:CONFig:CW | CW [A or B] | Enable CW mode on the specified input. |
| SENSe<sensor 1 or 2>:CONFig:MAP[:SYNChronized] | MAP [A or B] | Enable synchronised MAP mode on the specified input. |
| SENSe<sensor 1 or 2>:CONFig:MAP:UNSYnchronized | MAP [AorB] 0 | Enable unsynchronised MAP mode on the specified input |
| SENSe<sensor 1 or 2>:CONFig:BAP | BAP [A or B] | Enable BAP mode on the specified input. This command does not affect the burst start or end exclude or dropout tolerance parameters. The default values, set by *RST, correspond to the “automatic” settings in manual mode. |
| SENSe<sensor 1 or 2>:CONFig:PAP | PAP [A or B] | Enable PAP mode on the specified input. |
| SENSe<sensor 1 or 2>:CONFig:PAP:DCYCle space <Duty Cycle Value> (Where <Duty Cycle Value> is in the range 0.001 to 99.999% with a resolution of 0.001.) | DY | Set the duty cycle value that will be used when PAP is enabled. |
| SENSe<sensor 1 or 2>:CONFig:PAP:DCYCle? |  | Query the duty cycle |
| SENSe<sensor 1 or 2>:CONFig? | MEAS [A? or B?] | Query sensor configuration on specified input. Response is character response data: CW ¦ MAP ¦ MAPU ¦ BAP ¦ PAP.¦ UNCAL ¦ NOSENSOR |
| CALCulate<channel 1 to 4>:MODE space <NORMal, BURSt, SWIFt or FMOD selection> | FMOD | FMOD added to the list of available modes |
| SENSe<sensor 1 or 2>:GATE:MODE space <OFF, GATE, TRIGger or EDGE> Examples: SENS1:GATE:MODE GATE Enable external gating mode. SENS1:GATE:MODE? Read the current setting | GATE[A or B] (OFF, GATE, TRIGGER or EDGE) | Set up time gating mode on the specified input. OFF – Disable time gating. GATE – External gating mode. TRIG / TRIGGER – External trigger mode. EDGE – Burst edge mode. |
| SENSe<sensor 1 or 2>:GATE:POLarity space <INVert or NINVert) Examples: SENS1:GATE:POL NINV Set non-inverted polarity for the trigger input. SENS1:GATE:POL? Read the current setting. | GATE[A or B] (INVERT or NONINVERT) | Set the polarity of the trigger input signal. In external gating mode INVERT specifies that measurements can be taken when the trigger input is low. In external trigger mode, INVERT specifies that a falling edge on the trigger input triggers the gating. |
| SENSe<sensor 1 or 2>:GATE:DELay space <time> (Where <time> is in the range 0 to 327.675μs with a resolution of 5μs, where 0 represents some minimum non-zero delay time.) Examples: SENS1:GATE:DEL 20E-3 Set the gate delay to 20ms. SENS1:GATE:DEL? Read the current setting. The value is returned in units | GATE[A or B] DELAY <numeric value> | Set the delay time from the trigger input edge to the start of the gating period. |
| SENSe<sensor 1 or 2>:GATE:DURation space <time> (Where <time> is in the range 0 to 327.675μs with a resolution of 5μs.) Examples: SENS1:GATE:DUR 250E-3 Set the duration of the gating period to 250ms. SENS1:GATE:DUR? Read the current setting. | GATE[A or B] DURATION <numeric value> | Set the duration of the gating period.. This command is relevant only in the External Trigger Mode. |
| SENSe<sensor 1 or 2>:GATE:HOLDoff space <time> (Where <time> is in the range 0 to 327.675μs with a resolution of 5μs.) Examples: SENS1:GATE:HOLD 10E-3 Set the holdoff period to 10ms SENS1:GATE:HOLD? Read the current setting. | GATE[A or B] HOLDOFF <numeric value> | Set the hold-off time from the end of the gating period to the time when the circuit will accept another trigger input edge. This command is relevant only in the External Trigger mode. |
| CALCulate<channel 1 to 4>:SFUNctions:STATe space <On ¦ Off> Examples: CALC1:SFUN:STAT On Enable statistical functions. CALC1:SFUN:STAT? Read the current setting. | STAT [ON or OFF] | Enable / disable statistical functions (mean and standard deviation). |
| CALCulate<channel 1 to 4>:SFUNctions:DATa? | STAT? | Read the result produced by the statistics function. The response is in the form <mean>, <standard deviation> |