## calculator
Agilent Average Power Sensor Uncertainty Calculator |  |  | Instructions: Fill in all the blue-colored fields and the
| 2 |  | "Sensor Match" box; observe results in lower right.
Input | Value |  | Match Specification Style† | Distribution of ρ†
Device Under Test match (RL or SWR or ρ)* | 1.22 | VSWR | Maximum | C: Rayleigh
Sensor Model | E9300A |  | Maximum | C: Rayleigh
User-Entered Sensor Match* | 1.1 | VSWR | Maximum | C: Rayleigh
†These two columns set the form of the specification
and the PDF of the match. Hover below for details.
Power Meter Model | EPM
Frequency | 2 | GHz
Power | -13 | dBm |  | 5.0119e-05 | W
Number of Readings Averaged | 1
Average mode | Normal
Source of Uncertainty | Symbol | Value ± | Probability Distribution | Divisor | Result
Mismatch Gain Between Generator and Sensor | Mu | ¦Γmax¦g = 0.099 | C: Rayleigh | 4.1786 | 0.1129%
|  | ¦Γmax¦s = 0.048 | C: Rayleigh
Mismatch Gain Between Calibration Source and Sensor | Muc | ¦Г¦c = 0.029 | A: Uniform inside Circle | 2.4309 | 0.07313%
|  | ¦Гmax¦sc = 0.061 | C: Rayleigh
Power Meter Instrumentation Error | Pm | 0.5% | Gaussian | 2 | 0.25%
Power Meter Instrumentation Error During Calibration | Pmc | 0.5% | Gaussian | 2 | 0.25%
Power Meter Calibrator Output Power | Pcal | 0.5% | Gaussian | 2 | 0.25%
Zero Drift | D | 1.5e-10 | Gaussian | 2 | 0.0001496%
Power Sensor Calibration Factor Uncertainties | Kb | 1.7% | Gaussian | 2 | 0.85%
Power Sensor Linearity | Pl | 3% | Gaussian | 2 | 1.5%
Zero Set | Zs | 5e-10 | Gaussian | 2 | 0.0004988%
Sensor Noise | N | 7e-10 | Gaussian | 2 | 0.003841%
|  |  |  | Combined Uncertainty-RSSed = | 1.783%
|  |  |  | K = | 2
* This term can be entered as the return loss (RL) in dB (Sxx), or VSWR, or as a reflection coefficient. |  |  |  | Expanded Uncertainty = | 3.565%
e.g. RL = -15 dB is equivalent to VSWR = 1.43 is equivalent to ρ = 0.178 (reflection coefficient) |  |  |  | Upper Limit Uncertainty = | +0.152 dB
Enter 1 for VSWR = 1, a perfect match, or 0 for RL = 0 dB, a lossless reflection |  |  |  | Lower Limit Uncertainty = | −0.158 dB
Revision 9
Date updated: 5 March 2014
© Agilent Technologies, Inc. 2008, 2009, 2011 and 2014.
The data used in this calculator is based on the specifications in the EPM, EPM and P-Series datasheets.
The calculations are based on the ISO Guide to the Expression of Uncertainty in Measurement, often referred to as the GUM.
For more info about the GUM, and models used in this calculator, please refer to Agilent Application Note 1449-3,
"Fundamentals of RF and Microwave Power Measurements (Part 3)," literature number 5988-9215EN.
Note on D10: Match Specification Style: Set to 95th percentile or median if the specification is given in that format. "Maximum" is used for specifications given as warranted values. 80th percentile is useful for "typical" specifications when "typical" is defined as 80th percentile, such as in Agilent Signal Sources. "Typical" for Agilent Power Sensors is the Mean.
Note on E10: Distribution of ρ: The statistical distribution of the reflection coefficient, ρ, affects the uncertainty. In most cases, the pdf of ρ is a Rayleigh function, so this is the recommended setting for this column. Fixed setting is useful when the match is of known magnitude. Uniform inside a circle is useful when comparing with analyses using Case A of Agilent Application Note 1449-3. The letters A, B and C refer to the cases explained in that Application Note.
Sensor Match at Measurement Frequency: Data Source is
From Sensor Model (via Specifications vs. Frequency)
User-Entered; use the line above this box

## Data (hidden)
Power Meter instrumentation Errors & Calibrator Accuracy
|  |  | PowerMeter | Power Meter Instrumentation Error | Power Meter Instrumentation Error During Calibration | Power Meter Calibrator Output Power
|  |  | EPM | 0.5% | 0.5% | 0.5%
|  |  | EPM-P | 0.5% | 0.5% | 0.5%
|  |  | P-Series | 0.5% | 0.5% | 0.4%
Max SWR (25 +/-10C)
Frequency | 8481A | 8482A | 8483A | 8485A | 8487A | R8486A | Q8486A | 8481B | 8482B | 8481H | 8482H | 8481D | 8485D | 8487D | R8486D | Q8486D | E4412A | E4413A | E9300A | E9301A | E9304A | E9300B | E9301B | E9300H | E9301H | N8481A | N8482A | N8485A | N8481A-CFT | N8482A-CFT | N8485A-CFT | V8486A | W8486A | E8486A-100 | E8486A-200 | 8485A-033 | 8485D-033 | N8487A | N8488A | N8481B | N8482B | N8481H | N8482H | N8486AR | N8486AQ | N8487A-CFT | N8481B-CFT | N8482B-CFT | N8481H-CFT | N8482H-CFT | N8486AR-CFT | N8486AQ-CFT
9e-06 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.13 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
0.0001 | # | 1.6 | 1.8 | # | # | # | # | # | 1.1 | # | 1.2 | # | # | # | # | # | # | # | # | # | 1.13 | # | # | # | # | # | 1.52 | # | # | 1.52 | # | # | # | # | # | # | # | # | # | # | 1.08 | # | 1.13 | # | # | # | # | 1.08 | # | 1.13 | # | #
0.0003 | # | 1.2 | 1.8 | # | # | # | # | # | 1.1 | # | 1.2 | # | # | # | # | # | # | # | # | # | 1.13 | # | # | # | # | # | 1.12 | # | # | 1.12 | # | # | # | # | # | # | # | # | # | # | 1.08 | # | 1.13 | # | # | # | # | 1.08 | # | 1.13 | # | #
0.0006 | # | 1.2 | 1.18 | # | # | # | # | # | 1.1 | # | 1.2 | # | # | # | # | # | # | # | # | # | 1.13 | # | # | # | # | # | 1.12 | # | # | 1.12 | # | # | # | # | # | # | # | # | # | # | 1.08 | # | 1.13 | # | # | # | # | 1.08 | # | 1.13 | # | #
0.001 | # | 1.1 | 1.18 | # | # | # | # | # | 1.1 | # | 1.2 | # | # | # | # | # | # | # | # | # | 1.13 | # | # | # | # | # | 1.06 | # | # | 1.06 | # | # | # | # | # | # | # | # | # | # | 1.08 | # | 1.13 | # | # | # | # | 1.08 | # | 1.13 | # | #
0.01 | 1.4 | 1.1 | 1.18 | # | # | # | # | 1.1 | 1.1 | 1.2 | 1.2 | 1.4 | # | # | # | # | 1.22 | # | 1.15 | 1.15 | 1.13 | 1.12 | 1.12 | 1.15 | 1.15 | 1.37 | 1.06 | 1.33 | 1.37 | 1.06 | 1.33 | # | # | # | # | # | # | # | 1.03 | 1.09 | 1.08 | 1.16 | 1.13 | # | # | # | 1.09 | 1.08 | 1.16 | 1.13 | # | #
0.03 | 1.18 | 1.1 | 1.18 | # | # | # | # | 1.1 | 1.1 | 1.2 | 1.2 | 1.15 | # | # | # | # | 1.15 | # | 1.13 | 1.13 | 1.13 | 1.12 | 1.12 | 1.15 | 1.15 | 1.1 | 1.06 | 1.33 | 1.1 | 1.06 | 1.33 | # | # | # | # | # | # | # | 1.03 | 1.09 | 1.08 | 1.16 | 1.13 | # | # | # | 1.09 | 1.08 | 1.16 | 1.13 | # | #
0.05 | 1.1 | 1.1 | 1.18 | 1.15 | 1.15 | 1.15 | 1.15 | 1.1 | 1.1 | 1.2 | 1.2 | 1.15 | 1.19 | 1.19 | 1.4 | 1.4 | 1.15 | 1.21 | 1.13 | 1.13 | 1.13 | 1.12 | 1.12 | 1.15 | 1.15 | 1.05 | 1.06 | 1.06 | 1.05 | 1.06 | 1.06 | 1.073 | 1.076 | 1.066 | 1.077 | 1.15 | 1.19 | 1.08 | 1.03 | 1.09 | 1.08 | 1.16 | 1.13 | 1.17 | 1.17 | 1.08 | 1.09 | 1.08 | 1.16 | 1.13 | 1.17 | 1.17
0.1 | 1.1 | 1.1 | 1.18 | 1.1 | 1.1 | # | # | 1.1 | 1.1 | 1.2 | 1.2 | 1.15 | 1.15 | 1.15 | # | # | 1.15 | 1.19 | 1.13 | 1.13 | 1.13 | 1.12 | 1.12 | 1.15 | 1.15 | 1.05 | 1.06 | 1.04 | 1.05 | 1.06 | 1.04 | # | # | # | # | 1.1 | 1.15 | 1.05 | 1.04 | 1.09 | 1.08 | 1.16 | 1.13 | # | # | 1.05 | 1.09 | 1.08 | 1.16 | 1.13 | # | #
2 | 1.18 | 1.3 | 1.18 | 1.15 | 1.15 | # | # | 1.18 | 1.18 | 1.2 | 1.2 | 1.15 | 1.15 | 1.2 | # | # | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.12 | 1.15 | 1.15 | 1.16 | 1.07 | 1.13 | 1.16 | 1.07 | 1.13 | # | # | # | # | 1.15 | 1.15 | 1.1 | 1.04 | 1.14 | 1.16 | 1.16 | 1.13 | # | # | 1.1 | 1.14 | 1.16 | 1.16 | 1.13 | # | #
2.4 | 1.18 | 1.3 | 1.18 | 1.15 | 1.15 | # | # | 1.18 | 1.18 | 1.2 | 1.2 | 1.15 | 1.15 | 1.2 | # | # | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.12 | 1.15 | 1.15 | 1.16 | 1.07 | 1.13 | 1.16 | 1.07 | 1.13 | # | # | # | # | 1.15 | 1.15 | 1.1 | 1.09 | 1.14 | 1.16 | 1.16 | 1.13 | # | # | 1.1 | 1.14 | 1.16 | 1.16 | 1.13 | # | #
4 | 1.18 | 1.3 | # | 1.15 | 1.15 | # | # | 1.18 | 1.18 | 1.2 | 1.2 | 1.2 | 1.19 | 1.2 | # | # | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.12 | 1.15 | 1.15 | 1.16 | 1.07 | 1.13 | 1.16 | 1.07 | 1.13 | # | # | # | # | 1.15 | 1.19 | 1.1 | 1.09 | 1.14 | 1.16 | 1.16 | 1.13 | # | # | 1.1 | 1.14 | 1.16 | 1.16 | 1.13 | # | #
4.2 | 1.18 | 1.3 | # | 1.15 | 1.15 | # | # | 1.18 | 1.18 | 1.2 | 1.2 | 1.2 | 1.19 | 1.2 | # | # | 1.17 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.12 | 1.15 | 1.15 | 1.16 | 1.07 | 1.13 | 1.16 | 1.07 | 1.13 | # | # | # | # | 1.15 | 1.19 | 1.1 | 1.09 | 1.14 | 1.16 | 1.16 | 1.13 | # | # | 1.1 | 1.14 | 1.16 | 1.16 | 1.13 | # | #
6 | 1.18 | # | # | 1.15 | 1.15 | # | # | 1.18 | # | 1.2 | # | 1.2 | 1.19 | 1.2 | # | # | 1.2 | 1.19 | 1.19 | 1.19 | 1.19 | 1.12 | 1.12 | 1.15 | 1.15 | 1.16 | 1.07 | 1.13 | 1.16 | 1.07 | 1.13 | # | # | # | # | 1.15 | 1.19 | 1.1 | 1.09 | 1.14 | 1.16 | 1.16 | 1.13 | # | # | 1.1 | 1.14 | 1.16 | 1.16 | 1.13 | # | #
8 | 1.18 | # | # | 1.15 | 1.15 | # | # | 1.18 | # | 1.25 | # | 1.2 | 1.19 | 1.2 | # | # | 1.2 | 1.21 | 1.19 | # | # | 1.17 | # | 1.25 | # | 1.16 | # | 1.13 | 1.16 | # | 1.13 | # | # | # | # | 1.15 | 1.19 | 1.1 | 1.09 | 1.14 | # | 1.22 | # | # | # | 1.1 | 1.14 | # | 1.22 | # | # | #
10 | 1.18 | # | # | 1.15 | 1.15 | # | # | 1.18 | # | 1.25 | # | 1.3 | 1.19 | 1.2 | # | # | 1.2 | 1.21 | 1.19 | # | # | 1.17 | # | 1.25 | # | 1.16 | # | 1.13 | 1.16 | # | 1.13 | # | # | # | # | 1.15 | 1.19 | 1.1 | 1.09 | 1.14 | # | 1.22 | # | # | # | 1.1 | 1.14 | # | 1.22 | # | # | #
11 | 1.18 | # | # | 1.15 | 1.15 | # | # | 1.18 | # | 1.25 | # | 1.3 | 1.19 | 1.2 | # | # | 1.27 | 1.21 | 1.19 | # | # | 1.17 | # | 1.25 | # | 1.16 | # | 1.13 | 1.16 | # | 1.13 | # | # | # | # | 1.15 | 1.19 | 1.1 | 1.09 | 1.14 | # | 1.22 | # | # | # | 1.1 | 1.14 | # | 1.22 | # | # | #
12 | 1.18 | # | # | 1.15 | 1.15 | # | # | 1.18 | # | 1.25 | # | 1.3 | 1.25 | 1.2 | # | # | 1.27 | 1.21 | 1.19 | # | # | 1.17 | # | 1.25 | # | 1.16 | # | 1.13 | 1.16 | # | 1.13 | # | # | # | # | 1.15 | 1.25 | 1.1 | 1.09 | 1.14 | # | 1.22 | # | # | # | 1.1 | 1.14 | # | 1.22 | # | # | #
12.4 | 1.28 | # | # | 1.2 | 1.2 | # | # | 1.28 | # | 1.3 | # | 1.3 | 1.25 | 1.29 | # | # | 1.27 | 1.21 | 1.19 | # | # | 1.24 | # | 1.28 | # | 1.23 | # | 1.18 | 1.23 | # | 1.18 | # | # | # | # | 1.2 | 1.25 | 1.16 | 1.12 | 1.23 | # | 1.32 | # | # | # | 1.16 | 1.23 | # | 1.32 | # | # | #
14 | 1.28 | # | # | 1.2 | 1.2 | # | # | 1.28 | # | 1.3 | # | 1.3 | 1.25 | 1.29 | # | # | 1.27 | 1.21 | 1.22 | # | # | 1.24 | # | 1.28 | # | 1.23 | # | 1.18 | 1.23 | # | 1.18 | # | # | # | # | 1.2 | 1.25 | 1.16 | 1.12 | 1.23 | # | 1.32 | # | # | # | 1.16 | 1.23 | # | 1.32 | # | # | #
15 | 1.28 | # | # | 1.2 | 1.2 | # | # | 1.28 | # | 1.3 | # | 1.35 | 1.25 | 1.29 | # | # | 1.27 | 1.21 | 1.22 | # | # | 1.24 | # | 1.28 | # | 1.23 | # | 1.18 | 1.23 | # | 1.18 | # | # | # | # | 1.2 | 1.25 | 1.16 | 1.12 | 1.23 | # | 1.32 | # | # | # | 1.16 | 1.23 | # | 1.32 | # | # | #
16 | 1.28 | # | # | 1.2 | 1.2 | # | # | 1.28 | # | 1.3 | # | 1.35 | 1.25 | 1.29 | # | # | 1.27 | 1.21 | 1.26 | # | # | 1.24 | # | 1.28 | # | 1.23 | # | 1.18 | 1.23 | # | 1.18 | # | # | # | # | 1.2 | 1.25 | 1.16 | 1.12 | 1.23 | # | 1.32 | # | # | # | 1.16 | 1.23 | # | 1.32 | # | # | #
18 | 1.28 | # | # | 1.25 | 1.25 | # | # | 1.28 | # | 1.3 | # | 1.35 | 1.29 | 1.37 | # | # | 1.27 | 1.26 | 1.26 | # | # | 1.24 | # | 1.28 | # | 1.23 | # | 1.25 | 1.23 | # | 1.25 | # | # | # | # | 1.25 | 1.29 | 1.22 | 1.22 | 1.23 | # | 1.32 | # | # | # | 1.22 | 1.23 | # | 1.32 | # | # | #
18.1 | # | # | # | 1.25 | 1.25 | # | # | # | # | # | # | # | 1.29 | 1.37 | # | # | # | 1.26 | # | # | # | # | # | # | # | # | # | 1.25 | # | # | 1.25 | # | # | # | # | 1.25 | 1.29 | 1.22 | 1.22 | # | # | # | # | # | # | 1.22 | # | # | # | # | # | #
26.5 | # | # | # | 1.25 | 1.3 | 1.4 | # | # | # | # | # | # | 1.29 | 1.37 | 1.4 | # | # | 1.26 | # | # | # | # | # | # | # | # | # | 1.31 | # | # | 1.31 | # | # | # | # | 1.4 | 1.35 | 1.3 | 1.31 | # | # | # | # | 1.4 | # | 1.3 | # | # | # | # | 1.4 | #
26.6 | # | # | # | # | 1.3 | 1.4 | # | # | # | # | # | # | # | 1.37 | 1.4 | # | # | # | # | # | # | # | # | # | # | # | # | 1.31 | # | # | 1.31 | # | # | # | # | 1.4 | 1.35 | 1.3 | 1.31 | # | # | # | # | 1.4 | # | 1.3 | # | # | # | # | 1.4 | #
33 | # | # | # | # | 1.3 | 1.4 | 1.5 | # | # | # | # | # | # | 1.37 | 1.4 | 1.4 | # | # | # | # | # | # | # | # | # | # | # | 1.31 | # | # | 1.31 | # | # | # | # | 1.4 | 1.35 | 1.3 | 1.31 | # | # | # | # | 1.4 | 1.5 | 1.3 | # | # | # | # | 1.4 | 1.5
33.1 | # | # | # | # | 1.3 | 1.4 | 1.5 | # | # | # | # | # | # | 1.37 | 1.4 | 1.4 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.3 | 1.31 | # | # | # | # | 1.4 | 1.5 | 1.3 | # | # | # | # | 1.4 | 1.5
34 | # | # | # | # | 1.3 | 1.4 | 1.5 | # | # | # | # | # | # | 1.61 | 1.4 | 1.4 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.3 | 1.31 | # | # | # | # | 1.4 | 1.5 | 1.3 | # | # | # | # | 1.4 | 1.5
40 | # | # | # | # | 1.5 | 1.4 | 1.5 | # | # | # | # | # | # | 1.89 | 1.4 | 1.4 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.34 | 1.47 | # | # | # | # | 1.4 | 1.5 | 1.34 | # | # | # | # | 1.4 | 1.5
40.1 | # | # | # | # | 1.5 | # | 1.5 | # | # | # | # | # | # | 1.89 | # | 1.4 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.34 | 1.47 | # | # | # | # | # | 1.5 | 1.34 | # | # | # | # | # | 1.5
50 | # | # | # | # | 1.5 | # | 1.5 | # | # | # | # | # | # | 1.89 | # | 1.4 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.06 | # | # | # | # | # | 1.34 | 1.47 | # | # | # | # | # | 1.5 | 1.34 | # | # | # | # | # | 1.5
50.1 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.06 | # | # | # | # | # | # | 1.47 | # | # | # | # | # | # | # | # | # | # | # | # | #
60 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.06 | # | 1.06 | 1.28 | # | # | # | 1.47 | # | # | # | # | # | # | # | # | # | # | # | # | #
67 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.06 | # | 1.06 | 1.28 | # | # | # | 1.5 | # | # | # | # | # | # | # | # | # | # | # | # | #
70 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.06 | # | 1.06 | 1.28 | # | # | # | 1.5 | # | # | # | # | # | # | # | # | # | # | # | # | #
75 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.06 | 1.08 | 1.06 | 1.28 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
75.1 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.08 | 1.06 | 1.28 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
90 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.08 | 1.06 | 1.28 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
110 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.08 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
Power Linearity (25+/-10C)
Power | 8481A | 8482A | 8483A | 8485A | 8487A | R8486A | Q8486A | 8481B | 8482B | 8481H | 8482H | 8481D | 8485D | 8487D | R8486D | Q8486D | E4412A | E4413A | E9300A | E9301A | E9304A | E9300B | E9301B | E9300H | E9301H | N8481A | N8482A | N8485A | N8481A-CFT | N8482A-CFT | N8485A-CFT | V8486A | W8486A | E8486A-100 | E8486A-200 | 8485A-033 | 8485D-033 | N8487A | N8488A | N8481B | N8482B | N8481H | N8482H | N8486AR | N8486AQ | N8487A-CFT | N8481B-CFT | N8482B-CFT | N8481H-CFT | N8482H-CFT | N8486AR-CFT | N8486AQ-CFT
-70 | # | # | # | # | # | # | # | # | # | # | # | 0% | 0% | 0% | 0% | 0% | 3% | 3% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
-60 | # | # | # | # | # | # | # | # | # | # | # | 0% | 0% | 0% | 0% | 0% | 3% | 3% | 3% | 3% | 3% | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.5% | # | 0% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
-50 | # | # | # | # | # | # | # | # | # | # | # | 0% | 0% | 0% | 0% | 0% | 3% | 3% | 3% | 3% | 3% | # | # | 4% | 4% | # | # | # | # | # | # | # | # | # | 1.5% | # | 0% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
-35 | # | # | # | # | # | # | # | # | # | # | # | 0% | 0% | 0% | 0% | 0% | 3% | 3% | 3% | 3% | 3% | # | # | 4% | 4% | 0% | 0% | 0% | # | # | # | # | # | # | 1.5% | # | 0% | 0% | 0% | # | # | # | # | 0% | 0% | 0% | # | # | # | # | 0% | 0%
-30 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | # | # | 1% | 2% | 2% | 3% | 3% | 3% | 3% | 3% | 3% | 3% | 3.5% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0% | 0% | 0% | 1% | 2% | 1% | 1% | 0% | 2% | 0% | 0% | # | # | # | # | 0% | 0% | 0% | # | # | # | # | 0% | 0%
-25 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | # | # | 1% | 2% | 2% | 5% | 5% | 3% | 3% | 3% | 3% | 3% | 3.5% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0% | 0% | 0% | 1% | 2% | 1% | 1% | 0% | 2% | 0% | 0% | # | # | # | # | 0% | 0% | 0% | # | # | # | # | 0% | 0%
-20 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | # | # | 1% | 2% | 2% | 5% | 5% | 3% | 3% | 3% | 3% | 3% | 3.5% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0% | 0% | 0% | 1% | 2% | 1% | 1% | 0% | 2% | 0% | 0% | # | # | # | # | 0% | 0% | 0% | # | # | # | # | 0% | 0%
-15 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | # | # | 1% | 2% | 2% | 5% | 5% | 3% | 3% | 3% | 3% | 3% | 3.5% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0% | 0% | 0% | 1% | 2% | 1% | 1% | 0% | # | 0% | 0% | # | # | 0% | 0% | 0% | 0% | 0% | # | # | 0% | 0% | 0% | 0%
-10 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | 0% | 0% | # | # | # | # | # | 3% | 3% | 2.5% | 2.5% | 2.5% | 3.5% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0% | 0% | 0% | 1% | 2% | 1% | 1% | 0% | # | 0% | 0% | # | # | 0% | 0% | 0% | 0% | 0% | # | # | 0% | 0% | 0% | 0%
-5 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | 0% | 0% | # | # | # | # | # | 3% | 3% | 2.5% | 2.5% | 2.5% | 3.5% | 3.5% | 4% | 4% | 0% | 0% | 0% | 0% | 0% | 0% | 1% | 2% | 1% | 1% | 0% | # | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
-1 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | 0% | 0% | # | # | # | # | # | 3% | 3% | 2.5% | 2.5% | 2.5% | 3.5% | 3.5% | 4% | 4% | 0.52% | 0.52% | 0.52% | 0.52% | 0.52% | 0.52% | 1% | 2% | 1% | 1% | 0% | # | 0.52% | 0.52% | 0% | 0% | 0% | 0% | 0.52% | 0.52% | 0.52% | 0% | 0% | 0% | 0% | 0.52% | 0.52%
0 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | # | # | # | # | # | 3% | 3% | 2% | 2% | 2% | 3.5% | 3.5% | 3.5% | 3.5% | 0.52% | 0.52% | 0.52% | 0.52% | 0.52% | 0.52% | 1% | 2% | 1% | 1% | 0% | # | 0.52% | 0.52% | 0% | 0% | 0% | 0% | 0.52% | 0.52% | 0.52% | 0% | 0% | 0% | 0% | 0.52% | 0.52%
10 | 3% | 3% | 3% | 3% | 3% | 3% | 3% | 0% | 0% | 0% | 0% | # | # | # | # | # | 4.5% | 4.5% | 2% | 2% | 2% | 3.5% | 3.5% | 3% | 3% | 0.52% | 0.52% | 0.52% | 0.52% | 0.52% | 0.52% | 2% | 2% | 2% | 1.5% | 3% | # | 0.52% | 0.52% | 0% | 0% | 0% | 0% | 0.52% | 0.52% | 0.52% | 0% | 0% | 0% | 0% | 0.52% | 0.52%
15 | 3% | 3% | 3% | 3% | 3% | 3% | 3% | 0% | 0% | 0% | 0% | # | # | # | # | # | 4.5% | 4.5% | 2% | 2% | 2% | 3.5% | 3.5% | 3% | 3% | 0.8% | 0.8% | 0.8% | 0.8% | 0.8% | 0.8% | 2% | 2% | 2% | 1.5% | 3% | # | 0.8% | 0.8% | 0% | 0% | 0% | 0% | 0.8% | 0.8% | 0.8% | 0% | 0% | 0% | 0% | 0.8% | 0.8%
17 | 3% | 3% | 3% | 3% | 3% | 3% | 3% | 0% | 0% | 0% | 0% | # | # | # | # | # | 4.5% | 4.5% | 2% | 2% | 2% | 3.5% | 3.5% | 3% | 3% | 0.8% | 0.8% | 0.8% | 0.8% | 0.8% | 0.8% | 2% | 2% | 2% | 1.5% | 3% | # | 0.8% | 0.8% | 0% | 0% | 0.52% | 0.52% | 0.8% | 0.8% | 0.8% | 0% | 0% | 0.52% | 0.52% | 0.8% | 0.8%
20 | 3% | 3% | 3% | 3% | 3% | 3% | 3% | 0% | 0% | 0% | 0% | # | # | # | # | # | 4.5% | 4.5% | 2% | 2% | 2% | 3% | 3% | 3% | 3% | 0.8% | 0.8% | 0.8% | 0.8% | 0.8% | 0.8% | 2% | 2% | 2% | 1.5% | 3% | # | 0.8% | 0.8% | 0% | 0% | 0.52% | 0.52% | 0.8% | 0.8% | 0.8% | 0% | 0% | 0.52% | 0.52% | 0.8% | 0.8%
25 | # | # | # | # | # | # | # | 0% | 0% | 5% | 5% | # | # | # | # | # | # | # | # | # | # | 3% | 3% | 3% | 3% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0% | 0% | 0.52% | 0.52% | # | # | # | 0% | 0% | 0.52% | 0.52% | # | #
29 | # | # | # | # | # | # | # | 0% | 0% | 5% | 5% | # | # | # | # | # | # | # | # | # | # | 3% | 3% | 3% | 3% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0.52% | 0.52% | 0.52% | 0.52% | # | # | # | 0.52% | 0.52% | 0.52% | 0.52% | # | #
30 | # | # | # | # | # | # | # | 0% | 0% | 5% | 5% | # | # | # | # | # | # | # | # | # | # | 2.5% | 2.5% | 3% | 3% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0.52% | 0.52% | 0.8% | 0.8% | # | # | # | 0.52% | 0.52% | 0.8% | 0.8% | # | #
35 | # | # | # | # | # | # | # | 4% | 4% | 5% | 5% | # | # | # | # | # | # | # | # | # | # | 2.5% | 2.5% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0.52% | 0.52% | 0.8% | 0.8% | # | # | # | 0.52% | 0.52% | 0.8% | 0.8% | # | #
39 | # | # | # | # | # | # | # | 4% | 4% | # | # | # | # | # | # | # | # | # | # | # | # | 2.5% | 2.5% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0.8% | 0.8% | # | # | # | # | # | 0.8% | 0.8% | # | # | # | #
44 | # | # | # | # | # | # | # | 4% | 4% | # | # | # | # | # | # | # | # | # | # | # | # | 2.5% | 2.5% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0.8% | 0.8% | # | # | # | # | # | 0.8% | 0.8% | # | # | # | #
Lower Range | (Same specs for Upper/Lower for all models except for Meranti STD)
Power Range | Model | Zero Set | Zero Drift | Sensor noise
-30 to +20dBm | 8481A | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | 8482A | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | 8483A | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | 8485A | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | 8485A-033 | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | 8487A | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | R8486A | 5e-08 | 1e-08 | 1.1e-07
-30 to +20dBm | Q8486A | 5e-08 | 1e-08 | 1.1e-07
0 to +44dBm | 8481B | 5e-05 | 1e-05 | 0.00011
0 to +44dBm | 8482B | 5e-05 | 1e-05 | 0.00011
-10 to +35dBm | 8481H | 5e-06 | 1e-06 | 1e-05
-10 to +35dBm | 8482H | 5e-06 | 1e-06 | 1e-05
-70 to -20dBm | 8481D | 2e-11 | 4e-12 | 4.5e-11
-70 to -20dBm | 8485D | 2e-11 | 4e-12 | 4.5e-11
-70 to -20dBm | 8485D-033 | 2e-11 | 4e-12 | 4.5e-11
-70 to -20dBm | 8487D | 2e-11 | 4e-12 | 4.5e-11
-70 to -20dBm | R8486D | 3e-11 | 6e-12 | 6.5e-11
-70 to -20dBm | Q8486D | 3e-11 | 6e-12 | 6.5e-11
-30 to +20dBm | V8486A | 2e-07 | 4e-08 | 4.5e-07
-30 to +20dBm | W8486A | 2e-07 | 4e-08 | 4.5e-07
-70 to +20dBm | E4412A | 5e-11 | 1.5e-11 | 7e-11 |  | Lower range |  |  |  | Upper range
-70 to +20dBm | E4413A | 5e-11 | 1.5e-11 | 7e-11 |  | Model | Zero Set | Zero Drift | Sensor noise | Model | Zero Set | Zero Drift | Sensor noise
-60 to +20dBm | E9300A | 5e-10 | 1.5e-10 | 7e-10 |  | E9300A | 5e-10 | 1.5e-10 | 7e-10 | E9300A | 5e-07 | 1.5e-07 | 5e-07 | base on 15%-75% RH
-60 to +20dBm | E9301A | 5e-10 | 1.5e-10 | 7e-10 |  | E9301A | 5e-10 | 1.5e-10 | 7e-10 | E9301A | 5e-07 | 1.5e-07 | 5e-07 | base on 15%-75% RH
-60 to +20dBm | E9304A | 5e-10 | 1.5e-10 | 7e-10 |  | E9304A | 5e-10 | 1.5e-10 | 7e-10 | E9304A | 5e-07 | 1.5e-07 | 5e-07 | base on 15%-75% RH
-30 to +44dBm | E9300B | 5e-07 | 1.5e-07 | 7e-07 |  | E9300B | 5e-07 | 1.5e-07 | 7e-07 | E9300B | 0.0005 | 0.00015 | 0.0005 | base on 15%-75% RH
-30 to +44dBm | E9301B | 5e-07 | 1.5e-07 | 7e-07 |  | E9301B | 5e-07 | 1.5e-07 | 7e-07 | E9301B | 0.0005 | 0.00015 | 0.0005 | base on 15%-75% RH
-50 to +30dBm | E9300H | 5e-09 | 1.5e-09 | 7e-09 |  | E9300H | 5e-09 | 1.5e-09 | 7e-09 | E9300H | 5e-06 | 1.5e-06 | 5e-06 | base on 15%-75% RH
-50 to +30dBm | E9301H | 5e-09 | 1.5e-09 | 7e-09 |  | E9301H | 5e-09 | 1.5e-09 | 7e-09 | E9301H | 5e-06 | 1.5e-06 | 5e-06 | base on 15%-75% RH
-35 to -1dBm | N8481A | 2.5e-08 | 3e-09 | 8e-08 |  | N8481A | 2.5e-08 | 3e-09 | 8e-08 | N8481A | 6.3e-08 | 7e-09 | 1.14e-07 | base on 15%-75% RH
-35 to -1dBm | N8482A | 2.5e-08 | 3e-09 | 8e-08 |  | N8482A | 2.5e-08 | 3e-09 | 8e-08 | N8482A | 6.3e-08 | 7e-09 | 1.14e-07 | base on 15%-75% RH
-35 to -1dBm | N8485A | 2.5e-08 | 3e-09 | 8e-08 |  | N8485A | 2.5e-08 | 3e-09 | 8e-08 | N8485A | 6.3e-08 | 7e-09 | 1.14e-07 | base on 15%-75% RH
-30 to +20dBm | N8481A-CFT | 6.3e-08 | 7e-09 | 1.14e-07
-30 to +20dBm | N8482A-CFT | 6.3e-08 | 7e-09 | 1.14e-07
-30 to +20dBm | N8485A-CFT | 6.3e-08 | 7e-09 | 1.14e-07
-35 to -1dBm | N8487A | 2.5e-08 | 3e-09 | 8e-08 |  | N8487A | 2.5e-08 | 3e-09 | 8e-08 | N8487A | 6.3e-08 | 7e-09 | 1.14e-07
-35 to -1dBm | N8488A | 2.5e-08 | 3e-09 | 8e-08 |  | N8488A | 2.5e-08 | 3e-09 | 8e-08 | N8488A | 6.3e-08 | 7e-09 | 1.14e-07
-5 to +29dBm | N8481B | 2.5e-05 | 3e-06 | 8e-05 |  | N8481B | 2.5e-05 | 3e-06 | 8e-05 | N8481B | 6.3e-05 | 7e-06 | 0.000114
-5 to +29dBm | N8482B | 2.5e-05 | 3e-06 | 8e-05 |  | N8482B | 2.5e-05 | 3e-06 | 8e-05 | N8482B | 6.3e-05 | 7e-06 | 0.000114
-15 to +17dBm | N8481H | 2.5e-06 | 3e-07 | 8e-06 |  | N8481H | 2.5e-06 | 3e-07 | 8e-06 | N8481H | 6.3e-06 | 7e-07 | 1.14e-05
-15 to +17dBm | N8482H | 2.5e-06 | 3e-07 | 8e-06 |  | N8482H | 2.5e-06 | 3e-07 | 8e-06 | N8482H | 6.3e-06 | 7e-07 | 1.14e-05
-35 to -1dBm | N8486AR | 2.5e-08 | 3e-09 | 8e-08 |  | N8486AR | 2.5e-08 | 3e-09 | 8e-08 | N8486AR | 6.3e-08 | 7e-09 | 1.14e-07
-35 to -1dBm | N8486AQ | 2.5e-08 | 3e-09 | 8e-08 |  | N8486AQ | 2.5e-08 | 3e-09 | 8e-08 | N8486AQ | 6.3e-08 | 7e-09 | 1.14e-07
-30 to +20dBm | N8487A-CFT | 6.3e-08 | 7e-09 | 1.14e-07
0 to +44dBm | N8481B-CFT | 6.3e-05 | 7e-06 | 0.000114
0 to +44dBm | N8482B-CFT | 6.3e-05 | 7e-06 | 0.000114
-10 to +35dBm | N8481H-CFT | 6.3e-06 | 7e-07 | 1.14e-05
-10 to +35dBm | N8482H-CFT | 6.3e-06 | 7e-07 | 1.14e-05
-30 to +20dBm | N8486AR-CFT | 6.3e-08 | 7e-09 | 1.14e-07
-30 to +20dBm | N8486AQ-CFT | 6.3e-08 | 7e-09 | 1.14e-07
-30 to +20dBm | E8486A-100 | 2e-07 | 4e-08 | 4.5e-07
-60 to +20dBm | E8486A-200 | 2e-10 | 4e-11 | 4.5e-10
Normal | x2 | Normal(E4416/7A) | Average(N1911/2A)
Noise multiplier
Number of avg | Normal | x2 | Normal(E4416/7A) | Average(N1911/2A)
1 | 5.5 | 6.5 | 1 | 1
2 | 3.89 | 4.6 | 0.94 | 0.9
4 | 2.75 | 3.25 | 0.88 | 0.8
8 | 1.94 | 2.3 | 0.82 | 0.7
16 | 1 | 1.63 | 0.76 | 0.6
32 | 0.85 | 1 | 0.7 | 0.5
64 | 0.61 | 0.72 | 0.64 | 0.45
128 | 0.49 | 0.57 | 0.58 | 0.4
256 | 0.34 | 0.41 | 0.52 | 0.3
512 | 0.24 | 0.29 | 0.46 | 0.25
1024 | 0.17 | 0.2 | 0.4 | 0.2
Cal factor uncertainties (-35 to +20dBm, 25+/-3degC for Meranti and 8480, 25+/-10degC for E-Series)
Frequency | 8482A | 8482B | 8482H | 8483A | 8481A | 8481B | 8481H | 8481D | 8485A | 8485D | 8487A | 8487D | R8486A | Q8486A | R8486D | Q8486D | E4412A | E4413A | E9300A | E9301A | E9304A | E9300B | E9301B | E9300H | E9301H | N8481A | N8482A | N8485A | N8481A-CFT | N8482A-CFT | N8485A-CFT | V8486A | W8486A | E8486A-100 | E8486A-200 | 8485A-033 | 8485D-033 | N8487A | N8488A | N8481B | N8482B | N8481H | N8482H | N8486AR | N8486AQ | N8487A-CFT | N8481B-CFT | N8482B-CFT | N8481H-CFT | N8482H-CFT | N8486AR-CFT | N8486AQ-CFT
9e-06 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.6% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
0.0001 | 0.8% | 1.5% | 0.8% | 1.3% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.6% | # | # | # | # | # | 0.91% | # | # | 0.91% | # | # | # | # | # | # | # | # | # | # | 1.48% | # | 0.89% | # | # | # | # | 1.48% | # | 0.89% | # | #
0.0003 | 0.7% | 1.4% | 0.9% | 1.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.6% | # | # | # | # | # | 0.91% | # | # | 0.91% | # | # | # | # | # | # | # | # | # | # | 1.48% | # | 0.89% | # | # | # | # | 1.48% | # | 0.89% | # | #
0.001 | 0.7% | 1.4% | 0.9% | 1.1% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.6% | # | # | # | # | # | 0.91% | # | # | 0.91% | # | # | # | # | # | # | # | # | # | # | 1.48% | # | 0.89% | # | # | # | # | 1.48% | # | 0.89% | # | #
0.003 | 0.8% | 1.5% | 0.8% | 1.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1.6% | # | # | # | # | # | 0.91% | # | # | 0.91% | # | # | # | # | # | # | # | # | # | # | 1.48% | # | 0.89% | # | # | # | # | 1.48% | # | 0.89% | # | #
0.01 | 0.8% | 1.5% | 0.8% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | # | # | # | # | # | # | # | # | 1.8% | # | 1.8% | 1.8% | 1.6% | 1.8% | 1.8% | 1.8% | 1.8% | 0.82% | 0.78% | 0.82% | 0.82% | 0.78% | 0.82% | # | # | # | # | # | # | # | 1.7% | 1.42% | 1.43% | 0.77% | 0.79% | # | # | # | 1.42% | 1.43% | 0.77% | 0.79% | # | #
0.03 | 0.8% | 1.5% | 0.9% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | # | # | # | # | # | # | # | # | 1.8% | # | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 0.77% | 0.77% | 1.24% | 0.77% | 0.77% | 1.24% | # | # | # | # | # | # | # | 1.7% | 1.48% | 1.49% | 0.89% | 0.89% | # | # | # | 1.48% | 1.49% | 0.89% | 0.89% | # | #
0.05 | 0.7% | 1.4% | 0.8% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.38% | 1.3% | # | # | # | # | 0% | 0% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 0.77% | 0.77% | 1.24% | 0.77% | 0.77% | 1.24% | # | # | # | # | 1.4% | 1.4% | 1.33% | 1.6% | 1.48% | 1.49% | 0.89% | 0.89% | # | # | 1.33% | 1.48% | 1.49% | 0.89% | 0.89% | # | #
0.1 | 0.8% | 1.4% | 0.8% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.38% | 1.3% | # | # | # | # | 1.8% | 1.8% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 0.77% | 0.77% | 1.24% | 0.77% | 0.77% | 1.24% | # | # | # | # | 1.4% | 1.4% | 1.33% | 1.58% | 1.48% | 1.49% | 0.89% | 0.89% | # | # | 1.33% | 1.48% | 1.49% | 0.89% | 0.89% | # | #
0.3 | 0.8% | 1.4% | 0.8% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.38% | 1.3% | # | # | # | # | 1.8% | 1.8% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 1.6% | 0.77% | 0.77% | 1.24% | 0.77% | 0.77% | 1.24% | # | # | # | # | 1.4% | 1.4% | 1.33% | 1.58% | 1.48% | 1.49% | 0.89% | 0.89% | # | # | 1.33% | 1.48% | 1.49% | 0.89% | 0.89% | # | #
0.5 | 0.8% | 1.4% | 0.8% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.34% | 1.3% | # | # | # | # | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 0.78% | 0.78% | 1.26% | 0.78% | 0.78% | 1.26% | # | # | # | # | 1.4% | 1.4% | 1.35% | 1.58% | 1.48% | 1.49% | 0.89% | 0.89% | # | # | 1.35% | 1.48% | 1.49% | 0.89% | 0.89% | # | #
1 | 0.8% | 1.5% | 0.9% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.34% | 1.3% | # | # | # | # | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 0.78% | 0.78% | 1.26% | 0.78% | 0.78% | 1.26% | # | # | # | # | 1.4% | 1.4% | 1.35% | 1.58% | 1.48% | 1.49% | 0.89% | 0.89% | # | # | 1.35% | 1.48% | 1.49% | 0.89% | 0.89% | # | #
1.2 | 0.8% | 1.5% | 0.9% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.41% | 1.3% | # | # | # | # | 1.8% | 1.8% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 0.91% | 0.89% | 1.35% | 0.91% | 0.89% | 1.35% | # | # | # | # | 1.4% | 1.4% | 1.41% | 1.58% | 1.58% | 1.56% | 1.06% | 1.02% | # | # | 1.41% | 1.58% | 1.56% | 1.06% | 1.02% | # | #
2 | 0.8% | 1.5% | 0.8% | 1.2% | 0.7% | 1.4% | 0.8% | 0.8% | 1.4% | 1.4% | 1.41% | 1.3% | # | # | # | # | 2.4% | 2.4% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 0.91% | 0.89% | 1.35% | 0.91% | 0.89% | 1.35% | # | # | # | # | 1.4% | 1.4% | 1.41% | 1.75% | 1.58% | 1.56% | 1.06% | 1.02% | # | # | 1.41% | 1.58% | 1.56% | 1.06% | 1.02% | # | #
4 | 0.9% | 1.5% | 0.9% | # | 0.8% | 1.5% | 0.8% | 0.8% | 1.7% | 1.7% | 1.41% | 1.4% | # | # | # | # | 2.4% | 2.4% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 1.7% | 0.91% | 0.89% | 1.35% | 0.91% | 0.89% | 1.35% | # | # | # | # | 1.7% | 1.7% | 1.41% | 1.75% | 1.58% | 1.56% | 1.06% | 1.02% | # | # | 1.41% | 1.58% | 1.56% | 1.06% | 1.02% | # | #
6 | # | # | # | # | 0.9% | 1.5% | 0.9% | 0.9% | 1.7% | 1.7% | 1.59% | 1.4% | # | # | # | # | 2.4% | 2.4% | 1.8% | 1.7% | 1.7% | 1.8% | 1.8% | 1.8% | 1.7% | 1.26% | 0.89% | 1.66% | 1.26% | 0.89% | 1.66% | # | # | # | # | 1.7% | 1.7% | 1.61% | 1.75% | 1.77% | 1.56% | 1.46% | 1.02% | # | # | 1.61% | 1.77% | 1.56% | 1.46% | 1.02% | # | #
8 | # | # | # | # | 1% | 1.5% | 1% | 1% | 1.7% | 1.7% | 1.59% | 1.4% | # | # | # | # | 2.4% | 2.4% | 1.8% | # | # | 1.8% | # | 1.8% | # | 1.26% | # | 1.66% | 1.26% | # | 1.66% | # | # | # | # | 1.7% | 1.7% | 1.61% | 1.75% | 1.77% | # | 1.46% | # | # | # | 1.61% | 1.77% | # | 1.46% | # | # | #
10 | # | # | # | # | 1% | 1.6% | 1% | 1.1% | 1.9% | 1.9% | 1.59% | 1.5% | # | # | # | # | 2.4% | 2.6% | 1.8% | # | # | 1.8% | # | 1.8% | # | 1.26% | # | 1.66% | 1.26% | # | 1.66% | # | # | # | # | 1.9% | 1.9% | 1.61% | 1.75% | 1.77% | # | 1.46% | # | # | # | 1.61% | 1.77% | # | 1.46% | # | # | #
11 | # | # | # | # | 1% | 1.6% | 1% | 1.1% | 1.9% | 1.9% | 1.59% | 1.5% | # | # | # | # | 2.4% | 2.6% | 1.8% | # | # | 1.8% | # | 1.8% | # | 1.26% | # | 1.66% | 1.26% | # | 1.66% | # | # | # | # | 1.9% | 1.9% | 1.61% | 1.75% | 1.77% | # | 1.46% | # | # | # | 1.61% | 1.77% | # | 1.46% | # | # | #
12 | # | # | # | # | 1.1% | 1.6% | 1.1% | 1.2% | 1.9% | 1.9% | 1.59% | 1.5% | # | # | # | # | 2.4% | 2.8% | 1.8% | # | # | 1.8% | # | 1.8% | # | 1.26% | # | 1.66% | 1.26% | # | 1.66% | # | # | # | # | 1.9% | 1.9% | 1.61% | 1.75% | 1.77% | # | 1.46% | # | # | # | 1.61% | 1.77% | # | 1.46% | # | # | #
12.4 | # | # | # | # | 1.1% | 1.6% | 1.1% | 1.2% | 1.9% | 1.9% | 1.59% | 1.5% | # | # | # | # | 2.4% | 2.8% | 1.8% | # | # | 1.8% | # | 1.8% | # | 1.26% | # | 1.66% | 1.26% | # | 1.66% | # | # | # | # | 1.9% | 1.9% | 1.61% | 1.92% | 1.77% | # | 1.46% | # | # | # | 1.61% | 1.77% | # | 1.46% | # | # | #
14 | # | # | # | # | 1.2% | 1.6% | 1.2% | 1.1% | 2% | 2% | 1.69% | 1.6% | # | # | # | # | 2.4% | 2.8% | 2% | # | # | 2% | # | 2% | # | 1.59% | # | 1.83% | 1.59% | # | 1.83% | # | # | # | # | 2% | 2% | 1.73% | 1.92% | 1.92% | # | 1.73% | # | # | # | 1.73% | 1.92% | # | 1.73% | # | # | #
16 | # | # | # | # | 1.2% | 1.7% | 1.2% | 1.5% | 2% | 2.1% | 1.69% | 1.7% | # | # | # | # | 2.6% | 2.8% | 2% | # | # | 2% | # | 2% | # | 1.59% | # | 1.83% | 1.59% | # | 1.83% | # | # | # | # | 2% | 2.1% | 1.73% | 1.92% | 1.92% | # | 1.73% | # | # | # | 1.73% | 1.92% | # | 1.73% | # | # | #
17 | # | # | # | # | 1.2% | 1.7% | 1.2% | 1.5% | 2% | 2.1% | 1.69% | 1.7% | # | # | # | # | 2.6% | 2.8% | 2% | # | # | 2% | # | 2% | # | 1.59% | # | 1.83% | 1.59% | # | 1.83% | # | # | # | # | 2% | 2.1% | 1.73% | 1.92% | 1.92% | # | 1.73% | # | # | # | 1.73% | 1.92% | # | 1.73% | # | # | #
18 | # | # | # | # | 1.5% | 1.9% | 1.5% | 1.7% | 2.1% | 2.2% | 2.23% | 1.7% | # | # | # | # | 2.6% | 2.8% | 2% | # | # | 2% | # | 2% | # | 1.59% | # | 2.67% | 1.59% | # | 2.67% | # | # | # | # | 2.1% | 2.2% | 2.26% | 2.43% | 1.92% | # | 1.73% | # | # | # | 2.26% | 1.92% | # | 1.73% | # | # | #
20 | # | # | # | # | # | # | # | # | 2.1% | 2.7% | 2.23% | 1.7% | # | # | # | # | # | 3% | # | # | # | # | # | # | # | # | # | 2.67% | # | # | 2.67% | # | # | # | # | 2.1% | 2.7% | 2.26% | 2.43% | # | # | # | # | # | # | 2.26% | # | # | # | # | # | #
22 | # | # | # | # | # | # | # | # | 2.6% | 2.7% | 2.23% | 1.9% | # | # | # | # | # | 3% | # | # | # | # | # | # | # | # | # | 2.67% | # | # | 2.67% | # | # | # | # | 2.6% | 2.7% | 2.26% | 2.43% | # | # | # | # | # | # | 2.26% | # | # | # | # | # | #
24 | # | # | # | # | # | # | # | # | 2.6% | 2.7% | 2.23% | 1.9% | # | # | # | # | # | 3% | # | # | # | # | # | # | # | # | # | 2.67% | # | # | 2.67% | # | # | # | # | 2.6% | 2.7% | 2.26% | 2.43% | # | # | # | # | # | # | 2.26% | # | # | # | # | # | #
26 | # | # | # | # | # | # | # | # | 2.6% | 2.7% | 2.23% | 1.9% | # | # | # | # | # | 3% | # | # | # | # | # | # | # | # | # | 2.67% | # | # | 2.67% | # | # | # | # | 2.6% | 2.7% | 2.26% | 2.43% | # | # | # | # | # | # | 2.26% | # | # | # | # | # | #
26.5 | # | # | # | # | # | # | # | # | 2.8% | 2.8% | 2.58% | 2.2% | 2.2% | # | 3% | # | # | 3% | # | # | # | # | # | # | # | # | # | 3.32% | # | # | 3.32% | # | # | # | # | 2.8% | 2.8% | 2.58% | 4.56% | # | # | # | # | 2.63% | # | 2.58% | # | # | # | # | 2.63% | #
28 | # | # | # | # | # | # | # | # | 3.1% | 2.9% | 2.58% | 2.3% | 2.4% | # | 3.2% | # | # | 3% | # | # | # | # | # | # | # | # | # | 3.32% | # | # | 3.32% | # | # | # | # | 3.1% | 2.9% | 2.58% | 4.56% | # | # | # | # | 2.63% | # | 2.58% | # | # | # | # | 2.63% | #
30 | # | # | # | # | # | # | # | # | 3.2% | 3.2% | 2.58% | 2.4% | 2.5% | # | 3% | # | # | # | # | # | # | # | # | # | # | # | # | 3.32% | # | # | 3.32% | # | # | # | # | 3.2% | 3.2% | 2.58% | 4.56% | # | # | # | # | 2.63% | # | 2.58% | # | # | # | # | 2.63% | #
33 | # | # | # | # | # | # | # | # | 3.7% | 3.3% | 2.73% | 2.6% | 2.1% | 2.8% | 3% | 4.2% | # | # | # | # | # | # | # | # | # | # | # | 3.32% | # | # | 3.32% | # | # | # | # | 3.7% | 3.3% | 2.8% | 4.56% | # | # | # | # | 3.15% | 3.22% | 2.8% | # | # | # | # | 3.15% | 3.22%
34 | # | # | # | # | # | # | # | # | # | # | 2.73% | 2.6% | 2.1% | 2.8% | 3% | 4.2% | # | # | # | # | # | # | # | # | # | # | # | 103.3% | # | # | 103.3% | # | # | # | # | # | # | 2.8% | 4.56% | # | # | # | # | 3.15% | 3.46% | 2.8% | # | # | # | # | 3.15% | 3.46%
34.5 | # | # | # | # | # | # | # | # | # | # | 2.73% | 2.6% | 2.1% | 2.8% | 3% | 4.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2.8% | 4.56% | # | # | # | # | 3.15% | 3.46% | 2.8% | # | # | # | # | 3.15% | 3.46%
35 | # | # | # | # | # | # | # | # | # | # | 2.73% | 2.6% | 2.1% | 2.8% | 3% | 4.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2.8% | 4.56% | # | # | # | # | 3.15% | 3.22% | 2.8% | # | # | # | # | 3.15% | 3.22%
37 | # | # | # | # | # | # | # | # | # | # | 2.73% | 2.7% | 2.2% | 2.8% | 3% | 4.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2.8% | 4.56% | # | # | # | # | 3.15% | 3.22% | 2.8% | # | # | # | # | 3.15% | 3.22%
40 | # | # | # | # | # | # | # | # | # | # | 3.67% | 3% | 2.2% | 2.9% | 3% | 4.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 3.66% | 4.56% | # | # | # | # | 3.15% | 3.26% | 3.66% | # | # | # | # | 3.15% | 3.26%
42 | # | # | # | # | # | # | # | # | # | # | 3.67% | 3.2% | # | 3.9% | # | 4.9% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 3.66% | 4.56% | # | # | # | # | # | 3.26% | 3.66% | # | # | # | # | # | 3.26%
44 | # | # | # | # | # | # | # | # | # | # | 3.67% | 3.5% | # | 3.9% | # | 5.1% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 3.66% | 4.56% | # | # | # | # | # | 3.26% | 3.66% | # | # | # | # | # | 3.26%
45 | # | # | # | # | # | # | # | # | # | # | 4.33% | 3.5% | # | 3.9% | # | 5.1% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.23% | 4.56% | # | # | # | # | # | 3.33% | 4.23% | # | # | # | # | # | 3.33%
46 | # | # | # | # | # | # | # | # | # | # | 4.33% | 3.8% | # | 3.9% | # | 5.5% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.23% | 4.56% | # | # | # | # | # | 3.33% | 4.23% | # | # | # | # | # | 3.33%
48 | # | # | # | # | # | # | # | # | # | # | 4.33% | 3.8% | # | 4.9% | # | 5.8% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.23% | 4.56% | # | # | # | # | # | 3.33% | 4.23% | # | # | # | # | # | 3.33%
50 | # | # | # | # | # | # | # | # | # | # | 4.33% | 5% | # | 5.3% | # | 6.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.8% | # | # | # | # | # | 4.23% | 5.3% | # | # | # | # | # | 3.33% | 4.23% | # | # | # | # | # | 3.33%
51 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.1% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
52 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.9% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
53 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.9% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
54 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.9% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
55 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.6% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
56 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.1% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
57 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.2% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
58 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.2% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
59 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.2% | # | # | # | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
60 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.7% | # | 6.4% | 6.4% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
61 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.2% | # | 6.4% | 6.4% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
62 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.1% | # | 5.8% | 5.8% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
63 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6% | # | 5.6% | 5.6% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
64 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6% | # | 5.6% | 5.6% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
65 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.5% | # | 5.6% | 5.6% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
66 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.6% | # | 5.6% | 5.6% | # | # | # | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | #
67 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.7% | # | 5.8% | 5.8% | # | # | # | 5.81% | # | # | # | # | # | # | # | # | # | # | # | # | #
68 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.7% | # | 5.9% | 5.9% | # | # | # | 5.81% | # | # | # | # | # | # | # | # | # | # | # | # | #
69 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.6% | # | 5.9% | 5.9% | # | # | # | 5.81% | # | # | # | # | # | # | # | # | # | # | # | # | #
70 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.4% | # | 5.7% | 5.7% | # | # | # | 5.81% | # | # | # | # | # | # | # | # | # | # | # | # | #
71 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.7% | # | 5.6% | 5.6% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
72 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6.8% | # | 5.5% | 5.5% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
73 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 7% | # | 5.6% | 5.6% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
74 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 7.3% | # | 5.8% | 5.8% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
75 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.1% | 5.5% | 6.6% | 6.6% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
77 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.4% | 5.6% | 5.6% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
79 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.2% | 5.3% | 5.3% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
90 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.2% | 6.1% | 6.1% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
110 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5.2% | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | #
match | 1.22
rho DUT | 0.099099
rho sensor | 0.086758 | 1.19
rho sensor at cal | 0.061033 | 1.13
rho calibrator | 0.029126
Multiplier | 5.5
match | 1.1
rho user sensor | 0.047619
Calibrator Mismatch | EPM | EPM-P | P-Series
SWR calibrator | 1.06 | 1.06 | 1.08
rho calibrator | 0.029126 | 0.029126 | 0.038462
Uniform | 1.7321
Gaussian | 2
GammaStyles | Annotation | RayConv-95th | WithinCircConv | Boolean | 2nd Boolean
Maximum | ¦Γmax¦ | 0.712 | 1 | 0 | 1
95th %ile | ¦Γ95¦ | 1 | 1.026 | 0 | 0
80th %ile | ¦Γ80¦ | 1.269 | 1.118 | 0 | 0
Median | ¦Γ50¦ | 2.079 | 1.4142 | 0 | 0
Mean | ¦Γmean¦ | 1.953 | 1.5 | 0 | 0
Fixed | ¦Γ¦ | 1 | 1 | 1 | 1
GammaDistrGen | GammaMult | MuMult | Product
C: Rayleigh | 0.712 | 0.57776 | 0.41137
B: Fixed | 1 | 1 | 1
A: Uniform inside Circle | 1 | 0.70711 | 0.70711
GammaDistrSensor | GammaMult | MuMult | Product | Boolean
C: Rayleigh | 0.712 | 0.57776 | 0.41137 | 1
B: Fixed | 1 | 1 | 1 | 0
A: Uniform inside Circle | 1 | 0.70711 | 0.70711 | 1
GammaDistr50MHz | GammaMult | MuMult | Product
C: Rayleigh | 0.712 | 0.57776 | 0.41137
B: Fixed | 1 | 1 | 1
A: Uniform inside Circle | 1 | 0.70711 | 0.70711
Sensor SWR Option
2
Frequency Check
8481A
8481B
8481D
8481H
8482A
8482B
8482H
8483A
8485A
8485A-033
8485D
8485D-033
8487A
8487D
E4412A
E4413A
E9300A
E9300B
E9300H
E9301A
E9301B
E9301H
E9304A
N8481A
N8481A-CFT
N8481B
N8481B-CFT
N8481H
N8481H-CFT
N8482A
N8482A-CFT
N8482B
N8482B-CFT
N8482H
N8482H-CFT
N8485A
N8485A-CFT
N8486AQ | Out of sensor freq range!
N8486AQ-CFT | Out of sensor freq range!
N8486AR | Out of sensor freq range!
N8486AR-CFT | Out of sensor freq range!
N8487A
N8488A
N8487A-CFT
Q8486A | Out of sensor freq range!
Q8486D | Out of sensor freq range!
R8486A | Out of sensor freq range!
R8486D | Out of sensor freq range!
V8486A | Out of sensor freq range!
W8486A | Out of sensor freq range!
E8486A-100 | Out of sensor freq range!
E8486A-200 | Out of sensor freq range!
Power Check
8481A
8481B | Out of sensor power range!
8481D | Out of sensor power range!
8481H | Out of sensor power range!
8482A
8482B | Out of sensor power range!
8482H | Out of sensor power range!
8483A
8485A
8485A-033
8485D | Out of sensor power range!
8485D-033 | Out of sensor power range!
8487A
8487D | Out of sensor power range!
E4412A
E4413A
E9300A
E9300B
E9300H
E9301A
E9301B
E9301H
E9304A
N8481A
N8481A-CFT
N8481B | Out of sensor power range!
N8481B-CFT | Out of sensor power range!
N8481H
N8481H-CFT | Out of sensor power range!
N8482A
N8482A-CFT
N8482B | Out of sensor power range!
N8482B-CFT | Out of sensor power range!
N8482H
N8482H-CFT | Out of sensor power range!
N8485A
N8485A-CFT
N8486AQ
N8486AQ-CFT
N8486AR
N8486AR-CFT
N8487A
N8488A
N8487A-CFT
Q8486A
Q8486D | Out of sensor power range!
R8486A
R8486D | Out of sensor power range!
V8486A
W8486A
E8486A-100
E8486A-200
Note on E4: absolute accuracy
Note on G4: 23+/-3degC
Note on G5: 23+/-3degC
Note on E6: 0.5% for CW measurements. 0.8% for peak measurements
Note on G6: 25+/-10degC
Note on D138: Not used. This is default mode for use with E9320 series for peak, avg, time-related measurements.
Note on E138: Not used. This is for use with N1921/22A sensor only.
Note on A248: 0 to 55degC
Note on B263: Joe Gorin: Factor by which Gamma is multiplied to give the Gamma for which the formulas apply. For example, the Rayleigh formula applies for 95th percentile Gamma.
Note on C263: Joe Gorin: Multiplier of the measurement uncertainty.
Note on B268: Joe Gorin: Factor by which Gamma is multiplied to give the Gamma for which the formulas apply. For example, the Rayleigh formula applies for 95th percentile Gamma.
Note on C268: Joe Gorin: Multiplier of the measurement uncertainty.
Note on B273: Joe Gorin: Factor by which Gamma is multiplied to give the Gamma for which the formulas apply. For example, the Rayleigh formula applies for 95th percentile Gamma.
Note on C273: Joe Gorin: Multiplier of the measurement uncertainty.

## Revision history
Revision History
2008-11-11 | Rev 1
New creation
2009-03-17 | Rev 2
1. Corrected typo error on zero drift spec for E9300/1H
2. Corrected E4412/13A cal factor uncertainties for above 0dBm.
2009-04-24 | Rev 3
1. Added in more models - N8487A, N8481/2B, N8481/2H and N8486AR/AQ.
2. Updated Zero Set, Drift and Measurement Noise Spec for N8480 series.
3. Updated N8480 cal factor uncertainty values base on 25+/-3degC.
2009-11-05 | Rev 4
1. Corrected missing data on V/W8486A, 8485A-033 and 8485D-033.
2011-02-23 | Rev 5
1. Allowed DUT to be specified by maximum, 95th percentile, mean, etc, SWR
2. Allowed analysis base on Rayleigh distribution, fixed value (uniform on the edge of a circle), or uniform inside a circle in the complex plane
3. Allowed manual (user) entry of SWR for sensor at measurement frequency
4. Corrected error in Mismatch Gain Between Calibration Source and Sensor.
5. Added in new model - N8488A
6. Updated error in E4412/13A cal factor uncertainty above 0dBm.
2011-04-04 | Rev 6
1. Corrected error on column A7-A10 whereby error message will occur when user tried to change the radio button selection.
2011-07-01 | Rev 7
1. Updated error on E9300/1H (zero drift, zero set, noise at high power path).
2011-11-22 | Rev8
1. Corrected error on N8481B (zero drift, zero set, noise at low power path).
2. Updated CF MU for 8487A
2014-03-05 | Rev9
1. Added in new models - E8486A-100 and E8486A-200