(These changes and additions will be made to the E1430A Software Support Reference, as part of adding DOS support for the C library.)

# 2.0  Software Installation

## 2.1  Supported Configurations

(add...)

The following environments are supported by the C Interface Library.

HP Model V382 Embedded Controller running HP-UX 
HP Series 700 MXI running HP-UX
HP E1485 Signal Processor running SPOS
RadiSys EPC-7 Embedded Controller running DOS.

## 2.1 .. 2.4

(as is except add "HP-UX" several places to distinguish it from DOS)

## 2.5  DOS Disk Contents

The DOS disk contains the following sub directories:

EXAMPLE	example program
	INCLUDE	C include files such as e1430.h
	LIB	Interface library
	SCPI	SCPI driver and download utility
	SRC	Source code for the library
	TEST1430	diagnostic test program

## 2.6  Instructions for Installing  DOS Files

The disk directories and files should be copied to your hard disk under the directory \E1430.  The DOS XCOPY command can be used to copy all files to your hard disk.  For example:

MKDIR  C:\E1430
	XCOPY  /S  A:*.*  C:\E1430

# 3.   Overive of SCPI Support

## 3.1  Compiled SCPI

(no changes)

## 3.2  Downloadable SCPI Driver (HP-UX)

(no changes, but add..)

When using the SCPI drivers and the Command Module, the E1430 Logical Address must be set to a value that is a multiple of 8.  This is necessary because the secondary HP-IB address will be the logical address divided by 8.  The E1430 can then be controlled by sending SCPI command over the HP-IB, using the secondary address.

## 3.3  Downloadable SCPI Driver (DOS)

A DOS Personal Computer can be used to download the SCPI driver into the Command module memory.  During the download, the DOS computer communicates with the Command Module over an RS-232 link.  Once the SCPI driver is installed, the HP-IB interface on the Command Module.

The easiest way to download drivers is to use the HP VXI Installation Consultant (HP VIC) program.  This Windows program aids in setting configuration switches and downloading SCPI drivers.  HP VIC will require the driver file,  E1430.DC, in the \SCPI directory.

When using the SCPI drivers and the Command Module, the E1430 Logical Address must be set to a value that is a multiple of 8.  This is necessary because the secondary HP-IB address will be the logical address divided by 8.  The E1430 iscontrolled by sending SCPI command over the HP-IB, using this secondary address.

To download the E1430 SCPI driver to a Command Module without HP VIC, run the VXIDLD.EXE program in the SCPI directory of DOS disk.

# 6.   C Interface Library

## 6.1 Overview

The E1430 interface library provides access to the E1430 from C-language programs.  The library is a set of functions that allow the user to program the register based E1430 at a higher level that register read/writes.

(...more,  in E1430 Software Support manual)

The library is designed to work in several environments.  Hewlett-Packard provides compiled versions of the library for these environments:

HP-UX embedded VXI (or MXI-connected) computer.
	E1485 Signal Processor downloadable
	MS-DOS embedded VXI computer with SICL support

Source code for the library is provided, so it may be possible to use the library in other environments.

## 6.2 .. 6.10   (no change)

## 6.11   For the MS-DOS Environment

Your C program must be linked with the E1430 library, LIB1430.LIB.  Since LIB1430 uses SICL (Standard Instrument Control Language) to access the E1430, you must also link a SICL library.

The version of LIB1430.LIB provided with the E1430 should only be used with Microsoft compiler tools and SICL libraries for the RadiSys EPC-7 embedded computer.  If your environment is different, the library should be re-compiled.

To link your object file (example.obj), using Microsoft tools, the command is something like this:

link example.obj ,,, lib1430.lib+ mssicl.lib+ epcmsc.lib+ llibce.lib ;

mssicl.lib and epcmsc.lib are provided by RadiSys.  For a complete example of compiling and linking, see e1430/example/example.mak.

## 6.12   On Line Man Pages

(same as in the manual, except add...)
On-line man pages are not provided on the DOS environment.

## 6.13   An Example Program

The following C program is an example of using the E1430 library functions.  This program can be found in the file e1430/example.example.c.

############################################################################

#

# File:        e1430/example/example.c

# Description: Example program for HP E1430

# Language:    C

# (C) Copyright 1994, Hewlett-Packard Company, all rights reserved.

#

############################################################################

#

# This program demonstrates how to acquire data from the HP E1430 module

# over the VXI bus.  The program uses functions in the

# "E1430 C Interface Library", and so must be linked with this library.

#

# The program can run on several computer platforms, including:

#      An embedded HP-UX VXI computer like the V/382

#      An external HP-UX computer with MXI connection to VXI

#      An embedded DOS VXI computer like the RadiSys EPC-7

#      A Digital Signal Processor VXI module such as the E1485

#

############################################################################

*/

#include "machType.h"

#ifdef E1485_SOURCE

#  include "spil_d.h"

#else

#  include <stdio.h>

#  include <stdlib.h>

#endif

#include "e1430.h"

#define E1430_LA     129        /* VXI logical address of HP E1430 module */

#define BLOCK_SIZE   1024       /* Measurement blocksize      */

FLOATSIZ32 dataBuffer[BLOCK_SIZE];

ULONGSIZ32 dataBytes = BLOCK_SIZE * sizeof( FLOATSIZ32 );

SHORTSIZ16 groupID;            /* ID used by HP E1430 libraries           */

/**************************************************************************

*

> setUpInputs()

*

* $Description: $

*    This routine does all the necessary setup on the E1430 input module.

*

* $Return:     (void) $

*

**************************************************************************/

void setUpInputs( void )

{

SHORTSIZ16 error, la = E1430_LA;

/* Initialize the E1430 libraries, and reset the module */

/* Enable (1) or disable (0) e1430 debug printfs */

e1430_debug_level(0);

error = e1430_init_io_driver();

if( error )

{

(void)printf( "e1430_init_io_driver() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

groupID = e1430_create_module_group( 1, &la );

if( groupID <= 0 )

{

(void)printf( "Can't create module group with module at la = %d, %s\n",

la, e1430_get_error_string() );

exit( -1 );

}

error = e1430_reset_module( groupID );

if( error )

{

(void)printf( "e1430_reset_module() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

/* Set the range and other analog front end parameters */

error = e1430_set_analog_input( groupID,

2.0, /* range in Vp */

E1430_COUPLING_DC,

E1430_ANTIALIAS_ON,

E1430_INPUT_HI_CONN,

E1430_INPUT_LO_FLOAT );

if( error )

{

(void)printf( "e1430_set_analog_input() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

/* Set the block size and other data format parameters           */

error = e1430_set_data_format( groupID,

E1430_DATA_TYPE_REAL,

E1430_DATA_SIZE_32,

E1430_BLOCK_MODE,

BLOCK_SIZE,

E1430_APPEND_STATUS_OFF );

if( error )

{

(void)printf( "e1430_set_data_format() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

/* Set the frequency span by enabling the decimation filter.  */

error = e1430_set_decimation_filter( groupID,

1,   /* filter level.  1= 2MHz Bandwidth */

E1430_DECIMATION_ON,

E1430_ONEPASS,

E1430_PASS_TAG_32 );

if( error )

{

(void)printf( "e1430_set_decimation_filter() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

/* Autozero the inputs to remove DC offset */

error = e1430_auto_zero( groupID );

if( error )

{

(void)printf( "e1430_auto_zero() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

/* Note:  After a reset or autozero, the E1430 requires 20 seconds

* to allow the ADC correction values to settle to specified

* accuracy.   After code development, a delay should be inserted here.

*/

return;

}

/************************* Main program block ****************************/

int main( void )

{

SHORTSIZ16 error, adcOvld, adcErr;

LONGSIZ32  actualCnt, i;

(void)printf("E1430 example program.\n" );

(void)printf("Setting up ...\n" );

setUpInputs( );

(void)e1430_display_module_state( groupID );

(void)printf("Begin to acquire data ...\n" );

error = e1430_arm_module( groupID );

if( error )

{

(void)printf( "e1430_arm_module() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

(void)printf("Wait for data, then transfer data ...\n" );

error = e1430_read_float32_data( E1430_LA, dataBuffer, dataBytes,

&adcOvld, &adcErr,    &actualCnt );

if( error )

{

(void)printf( "e1430_read_float32_data() error: %d, %s\n",

error, e1430_get_error_string() );

exit( -1 );

}

(void)printf("Time record samples (in Volts): \n");

for (i = 0; i < 40; i += 4)

{

(void)printf("%3d: %10.6f  %10.6f  %10.6f  %10.6f \n", (int)i,

(double)dataBuffer[i+0], (double)dataBuffer[i+1],

(double)dataBuffer[i+2], (double)dataBuffer[i+3] );

}

/* delete E1430 module groups to free memory */

error = e1430_delete_all_module_groups();

return( 0 );

}

Here are some important things to notice in the example program:

The file e1430.h is included.  This file defines the constants and function prototypes for the library.

The function setUpInputs() initializes and sets up the E1430.  In this function, the order of the first few library function calls is important.  E1430_init_io_driver() must be called first to initialize the library.  This function also initializes the SICL I/O system used by the library.

Next, e1430_create_module_group() is called.  This function is passed a list of logical addresses, and it returns a groupID handle.  GroupID is used to reference the module group when calling other library functions.

Other functions, such as e1430_set_analog_input(), and e1430_set_data_format() are called to set up the E1430.  The order of these functions is not important.

Returning to the main program, e1430_display_module_state() prints the setup state of the E1430.

E1430_arm_module() will cause data collection to begin since triggering has not been set up.

E1430_read_float32_data() reads a block of data fills an array of floats.  The data is scaled in volts.

See the manual pages for more information on each library function.