Keithley 197a Autoranging Microvolt DMM

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Keithley 197a Autoranging Microvolt DMM.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

AT-MXI

MC-MXI

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

-  if you are using GPIB-VXI, there must be at least one unused

device name in the handler.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error Codes:

Error codes are returned as the return value of each instrument driver

function. A program should examine this value after each call to an instrument

driver function to determine if any error occurred. Possible error codes and

their meanings are listed with the corresponding instrument driver function.

________________________________________________________________________________

How To Use This Document:

This document is intended to be used as a programming reference manual.

It describes each function in the

Keithley 197a Autoranging Microvolt DMM

instrument. The functions are listed in alphabetical order with a

description of the function, C syntax of the function, a description of each

parameter, and possible error codes.

________________________________________________________________________________

Function Tree Layout:

Keithley 197a Autoranging Microvolt DMMFunction Name:

Initialize                                  kei197a_init

Configure

Query Measurement Function               kei197a_queryMeasurementFunction

Set dB Mode                              kei197a_setdBMode

Set Measurement Range                    kei197a_setMeasurementRange

Set Relative Mode                        kei197a_setRelativeMode

Set Trigger Mode                         kei197a_setTriggerMode

Send Trigger                                kei197a_sendTrigger

Query All Functions                         kei197a_queryAllFunctions

Measure                                     kei197a_measure

Close                                       kei197a_close

________________________________________________________________________________

Keithley 197a Autoranging Microvolt DMM

Keithley 197a Autoranging Microvolt DMM

This instrument module provides GPIB programming

support  for  the Keithley 197a DMM.  The

module is divided into the following functions:

(1) Initialize

Initialize the 197 on the  GPIB.

(2) Configure

Setup  parameters such as range, relative mode,

trigger mode and dB mode.

(3) Send Trigger

Sends the appropriate trigger to the instrument.

(4) Query all Functions

Reports the state of:

Measuring Function

Range

Relative Mode

Trigger Mode

(5) Measure

Read a single measurement value from the

meter.

(6) Close

Take the 197a offline from the GPIB.

This driver does not currently make use of the

data logger functions in the 197a.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

kei197a_close

void kei197a_close (void);

Purpose

CLOSE

Clears the multimeter, places it in  local mode,

and takes it off-line.

The  instrument must be reinitialized to use it

again.

________________________________________________________________________________

kei197a_init

void kei197a_init (int GPIBAddress);

Purpose

INITIALIZE

Initializes  the 197a in  the  following ways:

- clears the instrument with a GPIB device

clear

- sets the instrument to a known state based

on the factory default settings.

- queries for the ID of the instrument

- reads the measuring function that is selected

on the front panel

- checks to see if the dB mode is enabled

Parameter List

GPIBAddress

Variable Type       int

The GPIB ADDRESS control determines what address

on the GPIB will be used to communicate with the

instrument.  This value must  correspond  to the

current internal setting of the device.

Variable Type : Integer

Valid Range : 0 to 30

________________________________________________________________________________

kei197a_measure

void kei197a_measure (double *measurement);

Purpose

Measure

Triggers a measurement and reports a reading.

Parameter List

measurement

Variable Type       double (passed by reference)

________________________________________________________________________________

kei197a_queryAllFunctions

void kei197a_queryAllFunctions (char functionName[], int *functionCode,

char range[], int *rangeCode,

char relativeMode[], int *relativeModeCode,

char triggerMode[], int *triggerCode);

Purpose

Query All Functions

Returns the state of the following by

reading the status word of the 197a:

- measuring function (and dB mode)

- measurement range

- trigger mode

- relative mode

The states are reported in both the DDC code

and an easy-to-read string format.

This also makes the driver aware of any

front-panel commands that have affected

the four parameters listed above.

Parameter List

functionName

Variable Type       char []

functionCode

Variable Type       int (passed by reference)

range

Variable Type       char []

rangeCode

Variable Type       int (passed by reference)

relativeMode

Variable Type       char []

relativeModeCode

Variable Type       int (passed by reference)

triggerMode

Variable Type       char []

triggerCode

Variable Type       int (passed by reference)

________________________________________________________________________________

kei197a_queryMeasurementFunction

void kei197a_queryMeasurementFunction (char functionName[], int *functionCode);

Purpose

Query Measurement Function

- Returns the value of the measurement

function currently being performed

in both a string and DDC code format

Parameter List

functionName

Variable Type       char []

functionCode

Variable Type       int (passed by reference)

________________________________________________________________________________

kei197a_sendTrigger

void kei197a_sendTrigger (int trigger);

Purpose

Send Trigger

-     Sends Group Execute Trigger (GET) or the

the character 'X'

Parameter List

trigger

Variable Type       int

The TRIGGER control specifes a trigger to be

sent to the multimeter.

Variable Type : Integer

Valid Range:  0 - No action

1 - Send Group Execute Trigger

2 - Send the character X

________________________________________________________________________________

kei197a_setdBMode

void kei197a_setdBMode (int enable);

Purpose

Set dB Mode

- Enables or disables dB Mode

NOTE: The instrument must already be set

in VOLTS (AC or DC) mode on the front panel.

Parameter List

enable

Variable Type       int

________________________________________________________________________________

kei197a_setMeasurementRange

void kei197a_setMeasurementRange (double range);

Purpose

Set Measurement Range

- Selects the appropriate measurement range

based on the parameter "range"

NOTE: Current ranges are NOT programmable

over the GPIB on the 197a.

Parameter List

range

Variable Type       double

________________________________________________________________________________

kei197a_setRelativeMode

void kei197a_setRelativeMode (int enable);

Purpose

Set Relative Mode

-  Enables or disables relative mode.

-  Baseline suppression value is acquired

when command is sent.

Parameter List

enable

Variable Type       int

________________________________________________________________________________

kei197a_setTriggerMode

void kei197a_setTriggerMode (int triggerMode);

Purpose

Set Trigger Mode

- Sets the appropriate trigger mode.

Parameter List

triggerMode

Variable Type       int