## Slide 1: Time Domain Characterisation of Time and Frequency Standards
An Overview.
Richard Percival
Quartzlock (UK) Ltd

## Slide 2: Introduction
Who am I ?
Reason for paper
Who are Quartzlock?
What do we do
Notes: I would like to introduce myself to you today. My name is Richard Percival and I am from a UK based company called Quartzlock. We are manufacturers of a wide range of time and frequency generation, distribution and measurement instrumenttation. We have in our product range, active and passive masers, lf tracking receivers, OCXO's, distribution amplifiers, frequency convertors, frequency and phase comparators, GPS disciplined oscillators and rubidium frequency standards. At the moment we are developing a new low cost primary reference clock for telecoms and a low cost time transfer module. We are a small company and unlike may time and frequency oriented companies, we have retained a very strong metrology product like. The reason for this paper was originally to try and convey a little bit about the role measurement plays in time and frequency, with an overview of the method and statistics used within the community. The talk will hopefully not be too technical. Only you can tell me if I have succeeded in my aim!!!!

## Slide 3: Definitions
Accuracy- ability of a standard to deliver the frequency specified (i,e the Nameplate Frequency. A indication of adjustment quality.
Stability- ‘statistical estimate of the frequency fluctuations of a signal over a given period of time. Inherent quality of the oscillator

## Slide 4: Time and Frequency Statistics I
Problem with classical statistics- frequency output not solely ‘white’ noise
White noise can be averaged- phase advances compensated by phase retardation
Classical deviation will diverge-addition of more data leads deviation blow-up

## Slide 5: Time and Frequency Statistics II
Types of noise in Frequency Standards-white phase, flicker phase, white frequency, flicker frequency and random walk frequency.
Phase Noise- cause internal to device
Frequency Noise-environmental cause

## Slide 6: Development of new Statistics
Basic Premise-filter out the different noise present in atomic frequency standards
First differences-removes long term ageing and drift associated terms. ‘digital filter in the frequency domain’
Second differences-successful removal of all noise processes> characterisation!

## Slide 7: Quartzlock Data Acquisition
Phase comparison- Quartzlock A7 phase comparator, ‘suitable’ reference source time interval counter and frequency stability software. Frequency outputs
1pps comparison-time interval counter, suitable reference, frequency stability software. Time outputs

## Slide 8: Preparation for measurement
Characterisation of Comparator -identical signal split into system leads to coherent inputs> measurement noise shown
Short-Term Stability-Allan Variance gives indication of performance
Long term drift-temperature related drift must be quantified.

## Slide 9: Allan Variance- Important Points
Minimum  determined by sampling time - may be limited by instrumentation
Maximum  determined by run length - Allan Variance is a statistical measure. Max should be ~ run length/10

## Slide 10: Measurement Protocol
Decision Time Sampling time (minimum 1s & maximum 3200s) Run length ( needs to be 10x max Allan Variance  required. ). Typical ~10000s Single/ multi channel ?- single channel offers high resolution but multi channel offers speed. Types of characterisation - range of statistics available

## Slide 11: Measurement Procedure
Precautions- measurement system, reference standard and DUT suitably warm, connections suit stability levels (N-Type for high stability), input levels correct (6-13dBm)
Attitude- combination of impartiality and knowledge of ‘what to expect’. Difficult!!!

## Slide 12: Stability Characterisation
Allan Variance- developed in 1966, reveals non-white noise, good for long term frequency stability, CCIR stability definition
Modified Allan Variance- white and flicker phase separation, useful for STS in quartz and active masers, good for short term
Time Variance- concentrates on time,, ideal for telecom networks and clocks. Recent.

## Slide 13: Telecommunications Statistics
Maximum Time Interval Error (MTIE)- measure of peak time deviation of clock. Sensitive to any outliers or transients
rms. Time Interval Error (TIErms)- standard deviation of fractional freq fluctuations multiplied by averaging time. (zero freq. Offset) Time deviation better for divergent noise

## Slide 14: Statistical Analysis Procedure I
Data saved as ASCII data
Transfer to analysis PC
Open data file
Convert raw phase data to frequency data
Scaling correction made (if needed)
Outliers detected and optionally removed
Too many unexplained outliers > data rerun

## Slide 15: Statistical Analysis Procedure II
Phase and Frequency data plotted/analysed.
Daily (and monthly if enough data) drift computed and optionally removed (careful)
Frequency Stability (AVAR, MVAR and TVAR) computed and plotted
Optionally computation of Hadamard, Total, MTIE and rms.TIE

## Slide 16: Statistical Analysis Procedure III
All results printed out
Appropriate calibration and conformance certification filled in.
Failure of units are investigated and quality standards applied to prevent shipping of faulty units

## Slide 17: Conclusion
Basic Theory of Time and Frequency characterisation
Outline of variety of statistics used
Specific company calibration methodology.
More information?